## HS-VOC (hidden)
16-00866-N1 |  |  | HS-VOC
| Sample Mass (g) | Compound | Retention Time | Response | Area > 10 % | Conc (ug) | Conc (ug/g) | ID Level | CAS
Sample 1 | 4.78 | Ethylene Oxide | 1.56 | 5252650 | TRUE | 2.7713 | 0.57978 | TIC | 75-21-8
|  | Unknown | 2.69 | 284523 | TRUE | 0.15012 | 0.031405 | U | N/A
|  | Cyclohexane | 5.1 | 522608 | TRUE | 0.27573 | 0.057684 | MPC | 110-82-7
|  | Ethylene Benzene | 11.59 | 263870 | TRUE | 0.13922 | 0.029125 | MPC | 100-41-4
|  | p-xylene | 11.8 | 202187 | TRUE | 0.10668 | 0.022317 | TIC | 106-42-3
|  | Decane | 14.17 | 196653 | TRUE | 0.10376 | 0.021706 | MPC | 124-18-5
|  | Undecane | 16.11 | 329712 | TRUE | 0.17396 | 0.036393 | MPC | 1120-21-4
|  | Dodecane | 17.88 | 406980 | TRUE | 0.21473 | 0.044922 | MPC | 629-50-5
|  | Tridecane | 19.52 | 284356 | TRUE | 0.15003 | 0.031387 | MPC | 629-50-5
Sample 2 | 4.79 | Ethylene Oxide | 1.46 | 5382570 | TRUE | 2.8399 | 0.59288 | TIC | 75-21-8
|  | Unknown | 2.68 | 247449 | TRUE | 0.13056 | 0.027256 | U | N/A
|  | Cyclohexane | 5.09 | 568846 | TRUE | 0.30013 | 0.062657 | MPC | 110-82-7
|  | Ethylene Benzene | 11.6 | 272592 | TRUE | 0.14382 | 0.030025 | MPC | 100-41-4
|  | p-xylene | 11.8 | 236796 | TRUE | 0.12494 | 0.026082 | TIC | 106-42-3
|  | Decane | 14.17 | 122330 | FALSE | 0.064542 | 0.013474 | MPC | 124-18-5
|  | Undecane | 16.11 | 330961 | TRUE | 0.17462 | 0.036455 | MPC | 1120-21-4
|  | Unknown | 16.92 | 253812 | TRUE | 0.13391 | 0.027957 | U | N/A
|  | Dodecane | 17.88 | 436297 | TRUE | 0.23019 | 0.048057 | MPC | 629-50-5
|  | Naphthalene | 18.6 | 502096 | TRUE | 0.26491 | 0.055305 | MPC | 91-20-3
|  | 1,2,3-trichloro benzene | 19 | 298301 | TRUE | 0.15739 | 0.032857 | MPC | 87-61-6
|  | Tridecane | 19.52 | 274919 | TRUE | 0.14505 | 0.030282 | TIC | 629-50-5
Sample 3 | 4.79 | Ethylene Oxide | 1.46 | 5526490 | TRUE | 2.9158 | 0.60873 | TIC | 75-21-8
|  | Unknown | 2.69 | 251825 | TRUE | 0.13286 | 0.027738 | U | N/A
|  | Cyclohexane | 5.09 | 612789 | TRUE | 0.32331 | 0.067497 | MPC | 110-82-7
|  | Ethylene Benzene | 11.6 | 283055 | TRUE | 0.14934 | 0.031178 | MPC | 100-41-4
|  | p-xylene | 11.8 | 171869 | FALSE | 0.090679 | 0.018931 | TIC | 106-42-3
|  | Decane | 14.17 | 191250 | TRUE | 0.1009 | 0.021066 | MPC | 124-18-5
|  | Undecane | 16.11 | 403632 | TRUE | 0.21296 | 0.044459 | MPC | 1120-21-4
|  | Dodecane | 17.88 | 501143 | TRUE | 0.26441 | 0.0552 | MPC | 629-50-5
|  | Tridecane | 19.52 | 342529 | TRUE | 0.18072 | 0.037729 | TIC | 629-50-5
|  |  | Area Internal | 1895353 | Run ID: | A0074
|  |  | 10% area internal | 189535.3
|  |  | Conc (ug) | 1
Conc (ug) = Conc Standard (ug) * Area Sample/Area Standard
Conc (ug/g) = Conc (ug)/Sample Mass (g)
Highlighted results reported.
MPC - >80%
TIC - <80%
Note on B14: sdoherty: Update formula in Column G

## VOC
16-00866-N1 |  | VOC
[7 empty] | Control | Reportable
|  |  |  | Instrument | Dilution | Corrected | Corrected | Not Reportable
|  | Compound | Ret Time | Result (ug/L) | Factor | Result (ug/L) | Result (ug/L) | (R/NR) | ID | CAS #
1% S.A. | Control | Diethyl Ether | 2.573 | 6.4 | 20 | 128 | 128 | NR | IC | 60-29-7
10 minute |  | Benzyl Alcohol | 17.794 | 370.35 | 20 | 7407 | 7407 | R | MPC | 100-51-6
|  | Alkene (C10-C15) | 20.589 | 17.55 | 20 | 351 | 351 | R | U | N/A
|  | Alkene (C10-C15) | 20.798 | 105.81 | 20 | 2116.2 | 2116.2 | R | U | N/A
|  | Branched Alkene | 20.886 | 78.48 | 20 | 1569.6 | 1569.6 | R | U | N/A
|  | 2-ethyl-1-Hexanol | 21.007 | 37.27 | 20 | 745.4 | 745.4 | R | TIC | 104-76-7
|  | Alkene (C10-C15) | 21.48 | 5.67 | 20 | 113.4 | 113.4 | NR | U | N/A
|  | Unknown | 21.744 | 88.17 | 20 | 1763.4 | 1763.4 | NR | U | N/A
| Syringe S1 | Diethyl Ether | 2.574 | 2.25 | 20 | 45 | -83 | NR | IC | 60-29-7
|  | Benzyl Alcohol | 17.794 | 335.04 | 20 | 6700.8 | -706.2 | NR | MPC | 100-51-6
|  | Alkene (C10-C15) | 20.589 | 24.63 | 20 | 492.6 | 141.6 | NR | U | N/A
|  | Alkene (C10-C15) | 20.787 | 146.63 | 20 | 2932.6 | 816.4 | NR | U | N/A
|  | Branched Alkene | 20.886 | 100.4 | 20 | 2008 | 438.4 | NR | U | N/A
|  | 2-ethyl-1-Hexanol | 21.007 | 49.03 | 20 | 980.6 | 235.2 | NR | TIC | 104-76-7
|  | Alkene (C10-C15) | 21.491 | 31.15 | 20 | 623 | 509.6 | NR | U | N/A
|  | Unknown | 21.756 | 112.37 | 20 | 2247.4 | 484 | NR | U | N/A
| Syringe S2 | Diethyl Ether | 2.574 | 2.01 | 20 | 40.2 | -87.8 | NR | IC | 60-29-7
|  | Benzyl Alcohol | 17.794 | 295.57 | 20 | 5911.4 | -1495.6 | NR | MPC | 100-51-6
|  | Alkene (C10-C15) | 20.589 | 22.58 | 20 | 451.6 | 100.6 | NR | U | N/A
|  | Alkene (C10-C15) | 20.787 | 132.48 | 20 | 2649.6 | 533.4 | NR | U | N/A
|  | Branched Alkene | 20.886 | 92.99 | 20 | 1859.8 | 290.2 | NR | U | N/A
|  | 2-ethyl-1-Hexanol | 21.007 | 46.53 | 20 | 930.6 | 185.2 | NR | TIC | 104-76-7
|  | Alkene (C10-C15) | 21.491 | 56.87 | 20 | 1137.4 | 1024 | NR | U | N/A
|  | Unknown | 21.756 | 114.49 | 20 | 2289.8 | 526.4 | NR | U | N/A
| Syringe S3 | Diethyl Ether | 2.573 | 2.58 | 20 | 51.6 | -76.4 | NR | IC | 60-29-7
|  | Benzyl Alcohol | 17.794 | 249.8 | 20 | 4996 | -2411 | NR | MPC | 100-51-6
|  | Alkene (C10-C15) | 20.589 | 15.07 | 20 | 301.4 | -49.6 | NR | U | N/A
|  | Alkene (C10-C15) | 20.787 | 88.99 | 20 | 1779.8 | -336.4 | NR | U | N/A
|  | Branched Alkene | 20.886 | 73.22 | 20 | 1464.4 | -105.2 | NR | U | N/A
|  | 2-ethyl-1-Hexanol | 21.007 | 32.98 | 20 | 659.6 | -85.8 | NR | TIC | 104-76-7
|  | Alkene (C10-C15) | 21.491 | 71.85 | 20 | 1437 | 1323.6 | NR | U | N/A
|  | Unknown | 21.755 | 117.61 | 20 | 2352.2 | 588.8 | NR | U | N/A
| Catheter/Syringe S1 | Diethyl Ether | 2.573 | 2.99 | 20 | 59.8 | -68.2 | NR | IC | 60-29-7
|  | Benzyl Alcohol | 17.794 | 243.78 | 20 | 4875.6 | -2531.4 | NR | MPC | 100-51-6
|  | Alkene (C10-C15) | 20.589 | 15.06 | 20 | 301.2 | -49.8 | NR | U | N/A
|  | Alkene (C10-C15) | 20.787 | 88.6 | 20 | 1772 | -344.2 | NR | U | N/A
|  | Branched Alkene | 20.886 | 70.6 | 20 | 1412 | -157.6 | NR | U | N/A
|  | 2-ethyl-1-Hexanol | 21.007 | 31.51 | 20 | 630.2 | -115.2 | NR | TIC | 104-76-7
|  | Alkene (C10-C15) | 21.491 | 71.99 | 20 | 1439.8 | 1326.4 | NR | U | N/A
|  | Unknown | 21.756 | 109.77 | 20 | 2195.4 | 432 | NR | U | N/A
| Catheter/Syringe S2 | Diethyl Ether | 2.573 | 3.06 | 20 | 61.2 | -66.8 | NR | IC | 60-29-7
|  | Benzyl Alcohol | 17.794 | 283.88 | 20 | 5677.6 | -1729.4 | NR | MPC | 100-51-6
|  | Alkene (C10-C15) | 20.589 | 15.55 | 20 | 311 | -40 | NR | U | N/A
|  | Alkene (C10-C15) | 20.787 | 91.28 | 20 | 1825.6 | -290.6 | NR | U | N/A
|  | Branched Alkene | 20.886 | 72.58 | 20 | 1451.6 | -118 | NR | U | N/A
|  | 2-ethyl-1-Hexanol | 21.007 | 32.47 | 20 | 649.4 | -96 | NR | TIC | 104-76-7
|  | Alkene (C10-C15) | 21.491 | 70.15 | 20 | 1403 | 1289.6 | NR | U | N/A
|  | Unknown | 21.755 | 105.66 | 20 | 2113.2 | 349.8 | NR | U | N/A
| Catheter/Syringe S3 | Diethyl Ether | 2.573 | 3.47 | 20 | 69.4 | -58.6 | NR | IC | 60-29-7
|  | Benzyl Alcohol | 17.794 | 303.07 | 20 | 6061.4 | -1345.6 | NR | MPC | 100-51-6
|  | Alkene (C10-C15) | 20.589 | 16.97 | 20 | 339.4 | -11.6 | NR | U | N/A
|  | Alkene (C10-C15) | 20.787 | 103.47 | 20 | 2069.4 | -46.8 | NR | U | N/A
|  | Branched Alkene | 20.886 | 47.87 | 20 | 957.4 | -612.2 | NR | U | N/A
|  | 2-ethyl-1-Hexanol | 21.007 | 35.54 | 20 | 710.8 | -34.6 | NR | TIC | 104-76-7
1% S.A. | Control | Diethyl Ether | 2.573 | 5.57 | 20 | 111.4 | 111.4 | NR | IC | 60-29-7
30 minute |  | Benzyl Alcohol | 17.794 | 268.4 | 20 | 5368 | 5368 | R | MPC | 100-51-6
|  | Alkene (C10-C15) | 20.589 | 19.72 | 20 | 394.4 | 394.4 | R | U | N/A
|  | Alkene (C10-C15) | 20.787 | 112.8 | 20 | 2256 | 2256 | R | U | N/A
|  | Branched Alkene | 20.886 | 83.25 | 20 | 1665 | 1665 | R | U | N/A
|  | 2-ethyl-1-Hexanol | 21.007 | 39.43 | 20 | 788.6 | 788.6 | R | TIC | 104-76-7
|  | Alkene (C10-C15) | 21.48 | 64.52 | 20 | 1290.4 | 1290.4 | NR | U | N/A
|  | Unknown | 21.755 | 106.6 | 20 | 2132 | 2132 | NR | U | N/A
| Syringe S1 | Diethyl Ether | 2.573 | 2.26 | 20 | 45.2 | -66.2 | NR | IC | 60-29-7
|  | Benzyl Alcohol | 17.794 | 301.69 | 20 | 6033.8 | 665.8 | NR | MPC | 100-51-6
|  | Alkene (C10-C15) | 20.589 | 16.51 | 20 | 330.2 | -64.2 | NR | U | N/A
|  | Alkene (C10-C15) | 20.787 | 99.88 | 20 | 1997.6 | -258.4 | NR | U | N/A
|  | Branched Alkene | 20.886 | 79.94 | 20 | 1598.8 | -66.2 | NR | U | N/A
|  | 2-ethyl-1-Hexanol | 21.007 | 35.33 | 20 | 706.6 | -82 | NR | TIC | 104-76-7
|  | Alkene (C10-C15) | 21.491 | 69.78 | 20 | 1395.6 | 105.2 | NR | U | N/A
|  | Unknown | 21.756 | 116.07 | 20 | 2321.4 | 189.4 | NR | U | N/A
| Syringe S2 | Diethyl Ether | 2.573 | 2 | 20 | 40 | -71.4 | NR | IC | 60-29-7
|  | Benzyl Alcohol | 17.794 | 267.65 | 20 | 5353 | -15 | NR | MPC | 100-51-6
|  | Alkene (C10-C15) | 20.589 | 16.74 | 20 | 334.8 | -59.6 | NR | U | N/A
|  | Alkene (C10-C15) | 20.787 | 100.75 | 20 | 2015 | -241 | NR | U | N/A
|  | Branched Alkene | 20.886 | 73.72 | 20 | 1474.4 | -190.6 | NR | U | N/A
|  | 2-ethyl-1-Hexanol | 21.007 | 35.35 | 20 | 707 | -81.6 | NR | TIC | 104-76-7
|  | Alkene (C10-C15) | 21.491 | 72.28 | 20 | 1445.6 | 155.2 | NR | U | N/A
|  | Unknown | 21.756 | 99.47 | 20 | 1989.4 | -142.6 | NR | U | N/A
| Syringe S3 | Diethyl Ether | 2.573 | 1.78 | 20 | 35.6 | -75.8 | NR | IC | 60-29-7
|  | Benzyl Alcohol | 17.794 | 296.05 | 20 | 5921 | 553 | NR | MPC | 100-51-6
|  | Alkene (C10-C15) | 20.589 | 17.13 | 20 | 342.6 | -51.8 | NR | U | N/A
|  | Alkene (C10-C15) | 20.787 | 102.19 | 20 | 2043.8 | -212.2 | NR | U | N/A
|  | Branched Alkene | 20.886 | 74.38 | 20 | 1487.6 | -177.4 | NR | U | N/A
|  | 2-ethyl-1-Hexanol | 21.007 | 35.54 | 20 | 710.8 | -77.8 | NR | TIC | 104-76-7
|  | Alkene (C10-C15) | 21.491 | 73.47 | 20 | 1469.4 | 179 | NR | U | N/A
|  | Unknown | 21.755 | 99.48 | 20 | 1989.6 | -142.4 | NR | U | N/A
| Catheter/Syringe S1 | Diethyl Ether | 2.573 | 2.32 | 20 | 46.4 | -65 | NR | IC | 60-29-7
|  | Benzyl Alcohol | 17.794 | 318.47 | 20 | 6369.4 | 1001.4 | NR | MPC | 100-51-6
|  | Alkene (C10-C15) | 20.589 | 18.17 | 20 | 363.4 | -31 | NR | U | N/A
|  | Alkene (C10-C15) | 20.787 | 107.02 | 20 | 2140.4 | -115.6 | NR | U | N/A
|  | Branched Alkene | 20.886 | 75.46 | 20 | 1509.2 | -155.8 | NR | U | N/A
|  | 2-ethyl-1-Hexanol | 21.007 | 37.28 | 20 | 745.6 | -43 | NR | TIC | 104-76-7
|  | Alkene (C10-C15) | 21.491 | 71.35 | 20 | 1427 | 136.6 | NR | U | N/A
|  | Unknown | 21.756 | 94.38 | 20 | 1887.6 | -244.4 | NR | U | N/A
| Catheter/Syringe S2 | Diethyl Ether | 2.574 | 1.99 | 20 | 39.8 | -71.6 | NR | IC | 60-29-7
|  | Benzyl Alcohol | 17.794 | 296.9 | 20 | 5938 | 570 | NR | MPC | 100-51-6
|  | Alkene (C10-C15) | 20.589 | 19.29 | 20 | 385.8 | -8.6 | NR | U | N/A
|  | Alkene (C10-C15) | 20.798 | 108.06 | 20 | 2161.2 | -94.8 | NR | U | N/A
|  | Branched Alkene | 20.886 | 78.41 | 20 | 1568.2 | -96.8 | NR | U | N/A
|  | 2-ethyl-1-Hexanol | 21.007 | 37.85 | 20 | 757 | -31.6 | NR | TIC | 104-76-7
|  | Alkene (C10-C15) | 21.491 | 69.8 | 20 | 1396 | 105.6 | NR | U | N/A
|  | Unknown | 21.756 | 95.43 | 20 | 1908.6 | -223.4 | NR | U | N/A
| Catheter/Syringe S3 | Diethyl Ether | 2.573 | 2.17 | 20 | 43.4 | -68 | NR | IC | 60-29-7
|  | Benzyl Alcohol | 17.794 | 299 | 20 | 5980 | 612 | NR | MPC | 100-51-6
|  | Alkene (C10-C15) | 20.589 | 19.62 | 20 | 392.4 | -2 | NR | U | N/A
|  | Alkene (C10-C15) | 20.787 | 113.07 | 20 | 2261.4 | 5.4 | NR | U | N/A
|  | Branched Alkene | 20.886 | 80.33 | 20 | 1606.6 | -58.4 | NR | U | N/A
|  | 2-ethyl-1-Hexanol | 21.007 | 39.01 | 20 | 780.2 | -8.4 | NR | TIC | 104-76-7
|  | Alkene (C10-C15) | 21.491 | 63.74 | 20 | 1274.8 | -15.6 | NR | U | N/A
|  | Unknown | 21.756 | 94.74 | 20 | 1894.8 | -237.2 | NR | U | N/A
16-00866-N1 |  | VOC
[7 empty] | Control | Reportable
|  |  |  | Instrument | Dilution | Corrected | Corrected | Not Reportable
|  | Compound | Ret Time | Result (ug/L) | Factor | Result (ug/L) | Result (ug/L) | (R/NR) | ID | CAS #
3% S.A. | Control | Diethyl Ether | 2.574 | 12.55 | 20 | 251 | 251 | R | IC | 60-29-7
10 minute |  | Benzyl Alcohol | 17.794 | 305.65 | 20 | 6113 | 6113 | R | MPC | 100-51-6
|  | Alkene (C10-C15) | 20.457 | 8.68 | 20 | 173.6 | 173.6 | NR | U | N/A
|  | Alkene (C10-C15) | 20.589 | 76.18 | 20 | 1523.6 | 1523.6 | R | U | N/A
|  | Alkene (C10-C15) | 20.789 | 405.13 | 20 | 8102.6 | 8102.6 | R | U | N/A
|  | Branched Alkene | 20.886 | 178.02 | 20 | 3560.4 | 3560.4 | R | U | N/A
|  | 2-ethyl-1-Hexanol | 21.007 | 138.4 | 20 | 2768 | 2768 | R | TIC | 104-76-7
| Syringe S1 | Diethyl Ether | 2.574 | 5.85 | 20 | 117 | -134 | NR | IC | 60-29-7
|  | Benzyl Alcohol | 17.794 | 288.49 | 20 | 5769.8 | -343.2 | NR | MPC | 100-51-6
|  | (+/-)-3-Oxo-4a-methyl-trans-decalin | 19.929 | 5.15 | 20 | 103 | 103 | NR | TIC | 1197-95-1
|  | Alkene (C10-C15) | 20.457 | 10.46 | 20 | 209.2 | 209.2 | NR | U | N/A
|  | Alkene (C10-C15) | 20.589 | 71.9 | 20 | 1438 | -85.6 | NR | U | N/A
|  | Alkene (C10-C15) | 20.798 | 385.69 | 20 | 7713.8 | -388.8 | NR | U | N/A
|  | Branched Alkene | 20.886 | 254.41 | 20 | 5088.2 | 1527.8 | NR | U | N/A
|  | 2-ethyl-1-Hexanol | 21.007 | 145.44 | 20 | 2908.8 | 140.8 | NR | TIC | 104-76-7
|  | Alkene (C10-C15) | 21.491 | 56.61 | 20 | 1132.2 | 1132.2 | NR | U | N/A
|  | Unknown | 21.767 | 286 | 20 | 5720 | 5720 | NR | U | N/A
| Syringe S2 | Diethyl Ether | 2.573 | 5.09 | 20 | 101.8 | -149.2 | NR | IC | 60-29-7
|  | Benzyl Alcohol | 17.794 | 284.86 | 20 | 5697.2 | -415.8 | NR | MPC | 100-51-6
|  | (+/-)-3-Oxo-4a-methyl-trans-decalin | 19.918 | 5.75 | 20 | 115 | 115 | NR | TIC | 1197-95-1
|  | Alkene (C10-C15) | 20.457 | 10.67 | 20 | 213.4 | 213.4 | NR | U | N/A
|  | Alkene (C10-C15) | 20.589 | 70.38 | 20 | 1407.6 | -116 | NR | U | N/A
|  | Alkene (C10-C15) | 20.798 | 374.62 | 20 | 7492.4 | -610.2 | NR | U | N/A
|  | Branched Alkene | 20.886 | 255.37 | 20 | 5107.4 | 1547 | NR | U | N/A
|  | 2-ethyl-1-Hexanol | 21.007 | 137.25 | 20 | 2745 | -23 | NR | TIC | 104-76-7
|  | Alkene (C10-C15) | 21.491 | 150.29 | 20 | 3005.8 | 3005.8 | NR | U | N/A
|  | Unknown | 21.766 | 320.77 | 20 | 6415.4 | 6415.4 | NR | U | N/A
| Syringe S3 | Diethyl Ether | 2.573 | 5.16 | 20 | 103.2 | -147.8 | NR | IC | 60-29-7
|  | Benzyl Alcohol | 17.794 | 270.82 | 20 | 5416.4 | -696.6 | NR | MPC | 100-51-6
|  | (+/-)-3-Oxo-4a-methyl-trans-decalin | 19.918 | 6.47 | 20 | 129.4 | 129.4 | NR | TIC | 1197-95-1
|  | Alkene (C10-C15) | 20.457 | 11.6 | 20 | 232 | 232 | NR | U | N/A
|  | Alkene (C10-C15) | 20.589 | 76.84 | 20 | 1536.8 | 13.2 | NR | U | N/A
|  | Alkene (C10-C15) | 20.798 | 403.66 | 20 | 8073.2 | -29.4 | NR | U | N/A
|  | Branched Alkene | 20.886 | 274.93 | 20 | 5498.6 | 1938.2 | NR | U | N/A
|  | 2-ethyl-1-Hexanol | 21.007 | 148.98 | 20 | 2979.6 | 211.6 | NR | TIC | 104-76-7
|  | Alkene (C10-C15) | 21.491 | 161.82 | 20 | 3236.4 | 3236.4 | NR | U | N/A
|  | Unknown | 21.766 | 342.88 | 20 | 6857.6 | 6857.6 | NR | U | N/A
| Catheter/Syringe S1 | Diethyl Ether | 2.573 | 6.65 | 20 | 133 | -118 | NR | IC | 60-29-7
|  | Benzyl Alcohol | 17.794 | 264.8 | 20 | 5296 | -817 | NR | MPC | 100-51-6
|  | (+/-)-3-Oxo-4a-methyl-trans-decalin | 19.929 | 6.41 | 20 | 128.2 | 128.2 | NR | TIC | 1197-95-1
|  | Alkene (C10-C15) | 20.457 | 10.26 | 20 | 205.2 | 205.2 | NR | U | N/A
|  | Alkene (C10-C15) | 20.589 | 68.64 | 20 | 1372.8 | -150.8 | NR | U | N/A
|  | Alkene (C10-C15) | 20.798 | 358.49 | 20 | 7169.8 | -932.8 | NR | U | N/A
|  | Branched Alkene | 20.886 | 246.54 | 20 | 4930.8 | 1370.4 | NR | U | N/A
|  | 2-ethyl-1-Hexanol | 21.007 | 130.99 | 20 | 2619.8 | -148.2 | NR | TIC | 104-76-7
|  | Alkene (C10-C15) | 21.491 | 149.51 | 20 | 2990.2 | 2990.2 | NR | U | N/A
|  | Unknown | 21.767 | 320.98 | 20 | 6419.6 | 6419.6 | NR | U | N/A
| Catheter/Syringe S2 | Diethyl Ether | 2.573 | 6.98 | 20 | 139.6 | -111.4 | NR | IC | 60-29-7
|  | Benzyl Alcohol | 17.794 | 240.88 | 20 | 4817.6 | -1295.4 | NR | MPC | 100-51-6
|  | (+/-)-3-Oxo-4a-methyl-trans-decalin | 19.918 | 6.23 | 20 | 124.6 | 124.6 | NR | TIC | 1197-95-1
|  | Alkene (C10-C15) | 20.457 | 10.36 | 20 | 207.2 | 207.2 | NR | U | N/A
|  | Alkene (C10-C15) | 20.589 | 67.58 | 20 | 1351.6 | -172 | NR | U | N/A
|  | Alkene (C10-C15) | 20.798 | 355.63 | 20 | 7112.6 | -990 | NR | U | N/A
|  | Branched Alkene | 20.886 | 243.91 | 20 | 4878.2 | 1317.8 | NR | U | N/A
|  | 2-ethyl-1-Hexanol | 21.007 | 129.32 | 20 | 2586.4 | -181.6 | NR | TIC | 104-76-7
|  | Alkene (C10-C15) | 21.491 | 154.45 | 20 | 3089 | 3089 | NR | U | N/A
|  | Unknown | 21.766 | 303.11 | 20 | 6062.2 | 6062.2 | NR | U | N/A
| Catheter/Syringe S3 | Diethyl Ether | 2.573 | 5.16 | 20 | 103.2 | -147.8 | NR | IC | 60-29-7
|  | Benzyl Alcohol | 17.794 | 285.36 | 20 | 5707.2 | -405.8 | NR | MPC | 100-51-6
|  | (+/-)-3-Oxo-4a-methyl-trans-decalin | 19.918 | 5.88 | 20 | 117.6 | 117.6 | NR | TIC | 1197-95-1
|  | Alkene (C10-C15) | 20.457 | 11.47 | 20 | 229.4 | 229.4 | NR | U | N/A
|  | Alkene (C10-C15) | 20.589 | 74.51 | 20 | 1490.2 | -33.4 | NR | U | N/A
|  | Alkene (C10-C15) | 20.798 | 393.87 | 20 | 7877.4 | -225.2 | NR | U | N/A
|  | Branched Alkene | 20.886 | 254.97 | 20 | 5099.4 | 1539 | NR | U | N/A
|  | 2-ethyl-1-Hexanol | 21.007 | 142.7 | 20 | 2854 | 86 | NR | TIC | 104-76-7
|  | Alkene (C10-C15) | 21.491 | 154.91 | 20 | 3098.2 | 3098.2 | NR | U | N/A
|  | Unknown | 21.766 | 276.55 | 20 | 5531 | 5531 | NR | U | N/A
3% S.A. | Control | Diethyl Ether | 2.573 | 12.22 | 20 | 244.4 | 244.4 | R | IC | 60-29-7
30 minute |  | Benzyl Alcohol | 17.794 | 225.1 | 20 | 4502 | 4502 | R | MPC | 100-51-6
|  | Alkene (C10-C15) | 20.457 | 11.3 | 20 | 226 | 226 | R | U | N/A
|  | Alkene (C10-C15) | 20.589 | 92.3 | 20 | 1846 | 1846 | R | U | N/A
|  | Alkene (C10-C15) | 20.798 | 481 | 20 | 9620 | 9620 | R | U | N/A
|  | Branched Alkene | 20.886 | 213.1 | 20 | 4262 | 4262 | R | U | N/A
|  | 2-ethyl-1-Hexanol | 21.007 | 161.8 | 20 | 3236 | 3236 | R | TIC | 104-76-7
| Syringe S1 | Diethyl Ether | 2.573 | 2.88 | 20 | 57.6 | -186.8 | NR | IC | 60-29-7
|  | Benzyl Alcohol | 17.794 | 202.05 | 20 | 4041 | -461 | NR | MPC | 100-51-6
|  | (+/-)-3-Oxo-4a-methyl-trans-decalin | 19.918 | 4.62 | 20 | 92.4 | 92.4 | NR | TIC | 1197-95-1
|  | Alkene (C10-C15) | 20.457 | 5.48 | 20 | 109.6 | -116.4 | NR | U | N/A
|  | Alkene (C10-C15) | 20.589 | 42.96 | 20 | 859.2 | -986.8 | NR | U | N/A
|  | Alkene (C10-C15) | 20.798 | 227.87 | 20 | 4557.4 | -5062.6 | NR | U | N/A
|  | Branched Alkene | 20.886 | 179 | 20 | 3580 | -682 | NR | U | N/A
|  | 2-ethyl-1-Hexanol | 21.007 | 95.03 | 20 | 1900.6 | -1335.4 | NR | TIC | 104-76-7
|  | Octahydronaphthalen-4a-ol | 21.403 | 2.59 | 20 | 51.8 | 51.8 | NR | TIC | 55693-34-0
|  | Alkene (C10-C15) | 21.491 | 21.491 | 20 | 429.82 | 429.82 | NR | U | N/A
|  | Unknown | 21.755 | 257.2 | 20 | 5144 | 5144 | NR | U | N/A
| Syringe S2 | Diethyl Ether | 2.573 | 2.46 | 20 | 49.2 | -195.2 | NR | IC | 60-29-7
|  | Benzyl Alcohol | 17.794 | 229.76 | 20 | 4595.2 | 93.2 | NR | MPC | 100-51-6
|  | (+/-)-3-Oxo-4a-methyl-trans-decalin | 19.918 | 4.79 | 20 | 95.8 | 95.8 | NR | TIC | 1197-95-1
|  | Alkene (C10-C15) | 20.457 | 5.68 | 20 | 113.6 | -112.4 | NR | U | N/A
|  | Alkene (C10-C15) | 20.589 | 41.69 | 20 | 833.8 | -1012.2 | NR | U | N/A
|  | Alkene (C10-C15) | 20.798 | 222.91 | 20 | 4458.2 | -5161.8 | NR | U | N/A
|  | Branched Alkene | 20.886 | 181.72 | 20 | 3634.4 | -627.6 | NR | U | N/A
|  | 2-ethyl-1-Hexanol | 21.007 | 84.05 | 20 | 1681 | -1555 | NR | TIC | 104-76-7
|  | Alkene (C10-C15) | 21.491 | 111.19 | 20 | 2223.8 | 2223.8 | NR | U | N/A
|  | Unknown | 21.755 | 293.17 | 20 | 5863.4 | 5863.4 | NR | U | N/A
| Syringe S3 | Diethyl Ether | 2.574 | 2.27 | 20 | 45.4 | -199 | NR | IC | 60-29-7
|  | Benzyl Alcohol | 17.794 | 224.97 | 20 | 4499.4 | -2.6 | NR | MPC | 100-51-6
|  | (+/-)-3-Oxo-4a-methyl-trans-decalin | 19.918 | 5.13 | 20 | 102.6 | 102.6 | NR | TIC | 1197-95-1
|  | Alkene (C10-C15) | 20.457 | 6.2 | 20 | 124 | -102 | NR | U | N/A
|  | Alkene (C10-C15) | 20.589 | 43.96 | 20 | 879.2 | -966.8 | NR | U | N/A
|  | Alkene (C10-C15) | 20.798 | 237.88 | 20 | 4757.6 | -4862.4 | NR | U | N/A
|  | Branched Alkene | 20.886 | 187 | 20 | 3740 | -522 | NR | U | N/A
|  | 2-ethyl-1-Hexanol | 21.007 | 87.85 | 20 | 1757 | -1479 | NR | TIC | 104-76-7
|  | Alkene (C10-C15) | 21.491 | 115.1 | 20 | 2302 | 2302 | NR | U | N/A
|  | Unknown | 21.756 | 279.43 | 20 | 5588.6 | 5588.6 | NR | U | N/A
| Catheter/Syringe S1 | Diethyl Ether | 2.574 | 5.18 | 20 | 103.6 | -140.8 | NR | IC | 60-29-7
|  | Benzyl Alcohol | 17.794 | 247.2 | 20 | 4944 | 442 | NR | MPC | 100-51-6
|  | (+/-)-3-Oxo-4a-methyl-trans-decalin | 19.918 | 5.79 | 20 | 115.8 | 115.8 | NR | TIC | 1197-95-1
|  | Alkene (C10-C15) | 20.457 | 6.6 | 20 | 132 | -94 | NR | U | N/A
|  | Alkene (C10-C15) | 20.589 | 43.34 | 20 | 866.8 | -979.2 | NR | U | N/A
|  | Alkene (C10-C15) | 20.798 | 229.99 | 20 | 4599.8 | -5020.2 | NR | U | N/A
|  | Branched Alkene | 20.886 | 184.44 | 20 | 3688.8 | -573.2 | NR | U | N/A
|  | 2-ethyl-1-Hexanol | 21.007 | 84.48 | 20 | 1689.6 | -1546.4 | NR | TIC | 104-76-7
|  | Alkene (C10-C15) | 21.491 | 117.88 | 20 | 2357.6 | 2357.6 | NR | U | N/A
|  | Unknown | 21.756 | 287.08 | 20 | 5741.6 | 5741.6 | NR | U | N/A
| Catheter/Syringe S2 | Diethyl Ether | 2.573 | 5.26 | 20 | 105.2 | -139.2 | NR | IC | 60-29-7
|  | Benzyl Alcohol | 17.794 | 239.19 | 20 | 4783.8 | 281.8 | NR | MPC | 100-51-6
|  | (+/-)-3-Oxo-4a-methyl-trans-decalin | 19.918 | 5.4 | 20 | 108 | 108 | NR | TIC | 1197-95-1
|  | Alkene (C10-C15) | 20.457 | 7.1 | 20 | 142 | -84 | NR | U | N/A
|  | Alkene (C10-C15) | 20.589 | 46.97 | 20 | 939.4 | -906.6 | NR | U | N/A
|  | Alkene (C10-C15) | 20.798 | 250.23 | 20 | 5004.6 | -4615.4 | NR | U | N/A
|  | Branched Alkene | 20.886 | 189.38 | 20 | 3787.6 | -474.4 | NR | U | N/A
|  | 2-ethyl-1-Hexanol | 21.007 | 91.16 | 20 | 1823.2 | -1412.8 | NR | TIC | 104-76-7
|  | Alkene (C10-C15) | 21.491 | 125.34 | 20 | 2506.8 | 2506.8 | NR | U | N/A
|  | Unknown | 21.755 | 265.03 | 20 | 5300.6 | 5300.6 | NR | U | N/A
| Catheter/Syringe S3 | Diethyl Ether | 2.573 | 5.14 | 20 | 102.8 | -141.6 | NR | IC | 60-29-7
|  | Benzyl Alcohol | 17.794 | 293.8 | 20 | 5876 | 1374 | NR | MPC | 100-51-6
|  | (+/-)-3-Oxo-4a-methyl-trans-decalin | 19.918 | 5.71 | 20 | 114.2 | 114.2 | NR | TIC | 1197-95-1
|  | Alkene (C10-C15) | 20.457 | 8.02 | 20 | 160.4 | -65.6 | NR | U | N/A
|  | Alkene (C10-C15) | 20.589 | 51.13 | 20 | 1022.6 | -823.4 | NR | U | N/A
|  | Alkene (C10-C15) | 20.798 | 279.79 | 20 | 5595.8 | -4024.2 | NR | U | N/A
|  | Branched Alkene | 20.886 | 207.04 | 20 | 4140.8 | -121.2 | NR | U | N/A
|  | 2-ethyl-1-Hexanol | 21.007 | 101.04 | 20 | 2020.8 | -1215.2 | NR | TIC | 104-76-7
|  | Alkene (C10-C15) | 21.491 | 120.97 | 20 | 2419.4 | 2419.4 | NR | U | N/A
|  | Unknown | 21.755 | 279.3 | 20 | 5586 | 5586 | NR | U | N/A
Corrected Result (ug/L) = Instrument Result (ug/L) * Dilution Factor - Control Result (ug/L)
MPC - >80%
TIC - <80%

## SVOC
16-00866-N1 |  | SVOC
[8 empty] | Blank | Reportable
|  |  |  | Instrument | Concentration | Dilution | Corrected | Corrected | Not Reportable
|  | Compound | Ret Time | Result (ug/mL) | Factor | Factor | Result (ug/mL) | Result (ug/mL) | (R/NR) | ID | CAS #
1% S.A. | mb1603048OE | hexamethyl-Cyclotrisiloxane | 7.913 | 7.08 | 2 | 2 | 7.08 | 7.08 | NR | MPC | 541-05-9
10 minute |  | Benzyl Alcohol | 13.433 | 1993.48 | 2 | 2 | 1993.48 | 1993.48 | R | IC | 100-51-6
|  | Levoglucosenone | 14.931 | 12.14 | 2 | 2 | 12.14 | 12.14 | R | MPC | 37112-31-5
|  | C14 Alkene | 17.663 | 34.73 | 2 | 2 | 34.73 | 34.73 | R | U | N/A
|  | C14 Alkene | 17.788 | 124.91 | 2 | 2 | 124.91 | 124.91 | R | U | N/A
|  | C14 Alkene | 17.868 | 156.04 | 2 | 2 | 156.04 | 156.04 | R | U | N/A
|  | C14 Alkene | 18.165 | 375.65 | 2 | 2 | 375.65 | 375.65 | R | U | N/A
|  | C14 Alkane | 18.257 | 300.37 | 2 | 2 | 300.37 | 300.37 | R | U | N/A
|  | C14 Alkane | 18.348 | 250.33 | 2 | 2 | 250.33 | 250.33 | R | U | N/A
|  | Alkene | 18.485 | 12.46 | 2 | 2 | 12.46 | 12.46 | R | U | N/A
|  | Alkene | 18.577 | 14.62 | 2 | 2 | 14.62 | 14.62 | R | U | N/A
|  | Alkene | 18.657 | 16.3 | 2 | 2 | 16.3 | 16.3 | R | U | N/A
|  | Alkene | 18.702 | 15.87 | 2 | 2 | 15.87 | 15.87 | R | U | N/A
|  | Alkene | 18.771 | 41.11 | 2 | 2 | 41.11 | 41.11 | R | U | N/A
|  | Alkene | 18.931 | 147.16 | 2 | 2 | 147.16 | 147.16 | R | U | N/A
|  | Diphenylmethane | 21.011 | 14.85 | 2 | 2 | 14.85 | 14.85 | R | MPC | 101-81-5
|  | 7-ethyl-2-methyl-4-Undecanol | 21.4 | 565.71 | 2 | 2 | 565.71 | 565.71 | R | MPC | 103-20-8
|  | 1,1'-[oxybis(methylene)]bis-Benzene | 24.337 | 267.57 | 2 | 2 | 267.57 | 267.57 | R | MPC | 103-50-4
|  | 9,10-dihydro-Anthracene | 24.829 | 28.93 | 2 | 2 | 28.93 | 28.93 | R | MPC | 613-31-0
|  | Unknown | 26.337 | 7.89 | 2 | 2 | 7.89 | 7.89 | NR | U | N/A
|  | C16-C20 Alkyl Benzene | 27.572 | 8.37 | 2 | 2 | 8.37 | 8.37 | NR | U | N/A
|  | C16-C20 Alkyl Benzene | 27.778 | 70.49 | 2 | 2 | 70.49 | 70.49 | R | U | N/A
|  | Unknown | 28.932 | 4.85 | 2 | 2 | 4.85 | 4.85 | NR | U | N/A
|  | Unknown | 29.092 | 4.37 | 2 | 2 | 4.37 | 4.37 | NR | U | N/A
|  | Silane | 30.681 | 6.5 | 2 | 2 | 6.5 | 6.5 | NR | U | N/A
|  | 2-(phenylmethyl)-Benzenemethanol | 31.183 | 8.03 | 2 | 2 | 8.03 | 8.03 | NR | MPC | 1586-00-1
|  | Bibenzyl | 31.584 | 7.25 | 2 | 2 | 7.25 | 7.25 | NR | TIC | 103-29-7
|  | C20-C25 Phthalate | 31.767 | 8.68 | 2 | 2 | 8.68 | 8.68 | NR | U | N/A
|  | Unknown | 31.972 | 8.65 | 2 | 2 | 8.65 | 8.65 | NR | U | N/A
|  | Unknown | 32.681 | 9.89 | 2 | 2 | 9.89 | 9.89 | NR | U | N/A
|  | Unknown | 33.115 | 6.28 | 2 | 2 | 6.28 | 6.28 | NR | U | N/A
| Control | hexamethyl-Cyclotrisiloxane | 7.913 | 10.82 | 2 | 2 | 10.82 | 3.74 | NR | MPC | 541-05-9
|  | Benzyl Alcohol | 13.456 | 2195.17 | 2 | 2 | 2195.17 | 201.69 | NR | IC | 100-51-6
|  | Levoglucosenone | 14.931 | 14.68 | 2 | 2 | 14.68 | 2.54 | NR | MPC | 37112-31-5
|  | C14 Alkene | 17.662 | 34.08 | 2 | 2 | 34.08 | -0.65 | NR | U | N/A
|  | C14 Alkene | 17.788 | 124.21 | 2 | 2 | 124.21 | -0.7 | NR | U | N/A
|  | C14 Alkene | 17.868 | 159.27 | 2 | 2 | 159.27 | 3.23 | NR | U | N/A
|  | C14 Alkene | 18.165 | 376.49 | 2 | 2 | 376.49 | 0.84 | NR | U | N/A
|  | C14 Alkane | 18.257 | 302.9 | 2 | 2 | 302.9 | 2.53 | NR | U | N/A
|  | C14 Alkane | 18.348 | 255.8 | 2 | 2 | 255.8 | 5.47 | NR | U | N/A
|  | Alkene | 18.497 | 13.36 | 2 | 2 | 13.36 | 0.9 | NR | U | N/A
|  | Alkene | 18.577 | 15.98 | 2 | 2 | 15.98 | 1.36 | NR | U | N/A
|  | Alkene | 18.657 | 17.31 | 2 | 2 | 17.31 | 1.01 | NR | U | N/A
|  | Alkene | 18.703 | 16.65 | 2 | 2 | 16.65 | 0.78 | NR | U | N/A
|  | Alkene | 18.771 | 43.73 | 2 | 2 | 43.73 | 2.62 | NR | U | N/A
|  | Alkene | 18.931 | 148.84 | 2 | 2 | 148.84 | 1.68 | NR | U | N/A
|  | Diphenylmethane | 21.011 | 15.37 | 2 | 2 | 15.37 | 0.52 | NR | MPC | 101-81-5
|  | 7-ethyl-2-methyl-4-Undecanol | 21.4 | 810.95 | 2 | 2 | 810.95 | 245.24 | NR | MPC | 103-20-8
|  | 1,1'-[oxybis(methylene)]bis-Benzene | 24.337 | 552.6 | 2 | 2 | 552.6 | 285.03 | NR | MPC | 103-50-4
|  | 9,10-dihydro-Anthracene | 24.829 | 70.81 | 2 | 2 | 70.81 | 41.88 | NR | MPC | 613-31-0
|  | Unknown | 26.337 | 8.47 | 2 | 2 | 8.47 | 0.58 | NR | U | N/A
|  | C16-C20 Alkyl Benzene | 27.561 | 11.89 | 2 | 2 | 11.89 | 3.52 | NR | U | N/A
|  | C16-C20 Alkyl Benzene | 27.778 | 89.55 | 2 | 2 | 89.55 | 19.06 | NR | U | N/A
|  | Unknown | 28.921 | 8.04 | 2 | 2 | 8.04 | 3.19 | NR | U | N/A
|  | Unknown | 29.092 | 13.21 | 2 | 2 | 13.21 | 8.84 | NR | U | N/A
|  | Silane | 30.658 | 13.03 | 2 | 2 | 13.03 | 6.53 | NR | U | N/A
|  | 2-(phenylmethyl)-Benzenemethanol | 31.184 | 13.71 | 2 | 2 | 13.71 | 5.68 | NR | MPC | 1586-00-1
|  | Bibenzyl | 31.584 | 5.4 | 2 | 2 | 5.4 | -1.85 | NR | TIC | 103-29-7
|  | C20-C25 Phthalate | 31.767 | 19.31 | 2 | 2 | 19.31 | 10.63 | NR | U | N/A
|  | Unknown | 31.972 | 13.58 | 2 | 2 | 13.58 | 4.93 | NR | U | N/A
|  | Unknown | 32.121 | 10.95 | 2 | 2 | 10.95 | 10.95 | NR | U | N/A
|  | Unknown | 32.681 | 11.27 | 2 | 2 | 11.27 | 1.38 | NR | U | N/A
|  | Unknown | 33.104 | 24.34 | 2 | 2 | 24.34 | 18.06 | NR | U | N/A
| Catheter/Syringe S1 | Benzyl Alcohol | 13.479 | 2806.75 | 2 | 2 | 2806.75 | 813.27 | NR | IC | 100-51-6
|  | C14 Alkene | 17.651 | 18.17 | 2 | 2 | 18.17 | -16.56 | NR | U | N/A
|  | C14 Alkene | 17.788 | 70.05 | 2 | 2 | 70.05 | -54.86 | NR | U | N/A
|  | C14 Alkene | 17.868 | 115.96 | 2 | 2 | 115.96 | -40.08 | NR | U | N/A
|  | C14 Alkene | 18.165 | 295.79 | 2 | 2 | 295.79 | -79.86 | NR | U | N/A
|  | C14 Alkane | 18.257 | 229.05 | 2 | 2 | 229.05 | -71.32 | NR | U | N/A
|  | C14 Alkane | 18.348 | 184.76 | 2 | 2 | 184.76 | -65.57 | NR | U | N/A
|  | Alkene | 18.76 | 20.04 | 2 | 2 | 20.04 | 7.58 | NR | U | N/A
|  | Alkene | 18.931 | 76.97 | 2 | 2 | 76.97 | 62.35 | NR | U | N/A
|  | 7-ethyl-2-methyl-4-Undecanol | 21.411 | 727.73 | 2 | 2 | 727.73 | 162.02 | NR | MPC | 103-20-8
|  | 1,1'-[oxybis(methylene)]bis-Benzene | 24.337 | 278.46 | 2 | 2 | 278.46 | 10.89 | NR | MPC | 103-50-4
|  | 9,10-dihydro-Anthracene | 24.829 | 9.25 | 2 | 2 | 9.25 | -19.68 | NR | MPC | 613-31-0
|  | C16-C20 Alkyl Benzene | 27.572 | 8.39 | 2 | 2 | 8.39 | 0.02 | NR | U | N/A
|  | C16-C20 Alkyl Benzene | 27.778 | 76.59 | 2 | 2 | 76.59 | 6.1 | NR | U | N/A
|  | Unknown | 28.921 | 15.21 | 2 | 2 | 15.21 | 10.36 | NR | U | N/A
|  | Unknown | 29.104 | 28.08 | 2 | 2 | 28.08 | 23.71 | NR | U | N/A
|  | Silane | 30.658 | 18.62 | 2 | 2 | 18.62 | 12.12 | NR | U | N/A
|  | 2-(phenylmethyl)-Benzenemethanol | 31.184 | 12.87 | 2 | 2 | 12.87 | 4.84 | NR | MPC | 1586-00-1
|  | C20-C25 Phthalate | 31.767 | 26.8 | 2 | 2 | 26.8 | 18.12 | NR | U | N/A
|  | Unknown | 33.104 | 16.48 | 2 | 2 | 16.48 | 7.83 | NR | U | N/A
| Catheter/Syringe S2 | Benzyl Alcohol | 13.491 | 2722.18 | 2 | 2 | 2722.18 | 728.7 | NR | IC | 100-51-6
|  | C14 Alkene | 17.651 | 19.69 | 2 | 2 | 19.69 | -15.04 | NR | U | N/A
|  | C14 Alkene | 17.788 | 74.73 | 2 | 2 | 74.73 | -50.18 | NR | U | N/A
|  | C14 Alkene | 17.868 | 117.3 | 2 | 2 | 117.3 | -38.74 | NR | U | N/A
|  | C14 Alkene | 18.166 | 294.4 | 2 | 2 | 294.4 | -81.25 | NR | U | N/A
|  | C14 Alkane | 18.257 | 227.42 | 2 | 2 | 227.42 | -72.95 | NR | U | N/A
|  | C14 Alkane | 18.348 | 188.02 | 2 | 2 | 188.02 | -62.31 | NR | U | N/A
|  | Alkene | 18.76 | 26.69 | 2 | 2 | 26.69 | 14.23 | NR | U | N/A
|  | Alkene | 18.92 | 85.02 | 2 | 2 | 85.02 | 70.4 | NR | U | N/A
|  | 7-ethyl-2-methyl-4-Undecanol | 21.412 | 826.45 | 2 | 2 | 826.45 | 260.74 | NR | MPC | 103-20-8
|  | 1,1'-[oxybis(methylene)]bis-Benzene | 24.337 | 264.15 | 2 | 2 | 264.15 | -3.42 | NR | MPC | 103-50-4
|  | 9,10-dihydro-Anthracene | 24.829 | 11.64 | 2 | 2 | 11.64 | -17.29 | NR | MPC | 613-31-0
|  | C16-C20 Alkyl Benzene | 27.572 | 9.14 | 2 | 2 | 9.14 | 0.77 | NR | U | N/A
|  | C16-C20 Alkyl Benzene | 27.778 | 71.69 | 2 | 2 | 71.69 | 1.2 | NR | U | N/A
|  | Unknown | 28.92 | 18.21 | 2 | 2 | 18.21 | 13.36 | NR | U | N/A
|  | Unknown | 29.092 | 25.17 | 2 | 2 | 25.17 | 20.8 | NR | U | N/A
|  | Silane | 30.658 | 23.08 | 2 | 2 | 23.08 | 16.58 | NR | U | N/A
|  | 2-(phenylmethyl)-Benzenemethanol | 31.184 | 11.49 | 2 | 2 | 11.49 | 3.46 | NR | MPC | 1586-00-1
|  | C20-C25 Phthalate | 31.767 | 32.76 | 2 | 2 | 32.76 | 24.08 | NR | U | N/A
|  | Unknown | 33.093 | 20.8 | 2 | 2 | 20.8 | 12.15 | NR | U | N/A
| Catheter/Syringe S3 | Benzyl Alcohol | 13.468 | 2627.17 | 2 | 2 | 2627.17 | 633.69 | NR | IC | 100-51-6
|  | C14 Alkene | 17.651 | 17.29 | 2 | 2 | 17.29 | -17.44 | NR | U | N/A
|  | C14 Alkene | 17.788 | 66.81 | 2 | 2 | 66.81 | -58.1 | NR | U | N/A
|  | C14 Alkene | 17.868 | 100.52 | 2 | 2 | 100.52 | -55.52 | NR | U | N/A
|  | C14 Alkene | 18.165 | 261.12 | 2 | 2 | 261.12 | -114.53 | NR | U | N/A
|  | C14 Alkane | 18.245 | 205.95 | 2 | 2 | 205.95 | -94.42 | NR | U | N/A
|  | C14 Alkane | 18.348 | 166.54 | 2 | 2 | 166.54 | -83.79 | NR | U | N/A
|  | Alkene | 18.76 | 23.21 | 2 | 2 | 23.21 | 10.75 | NR | U | N/A
|  | Alkene | 18.92 | 77.58 | 2 | 2 | 77.58 | 62.96 | NR | U | N/A
|  | 7-ethyl-2-methyl-4-Undecanol | 21.4 | 769.75 | 2 | 2 | 769.75 | 204.04 | NR | MPC | 103-20-8
|  | 1,1'-[oxybis(methylene)]bis-Benzene | 24.337 | 247.56 | 2 | 2 | 247.56 | -20.01 | NR | MPC | 103-50-4
|  | 9,10-dihydro-Anthracene | 24.829 | 18.58 | 2 | 2 | 18.58 | -10.35 | NR | MPC | 613-31-0
|  | C16-C20 Alkyl Benzene | 27.572 | 9.3 | 2 | 2 | 9.3 | 0.93 | NR | U | N/A
|  | C16-C20 Alkyl Benzene | 27.778 | 68.56 | 2 | 2 | 68.56 | -1.93 | NR | U | N/A
|  | Unknown | 28.921 | 8.91 | 2 | 2 | 8.91 | 4.06 | NR | U | N/A
|  | Unknown | 29.092 | 23.49 | 2 | 2 | 23.49 | 19.12 | NR | U | N/A
|  | Silane | 30.658 | 20.54 | 2 | 2 | 20.54 | 14.04 | NR | U | N/A
|  | 2-(phenylmethyl)-Benzenemethanol | 31.184 | 13.82 | 2 | 2 | 13.82 | 5.79 | NR | MPC | 1586-00-1
|  | C20-C25 Phthalate | 31.767 | 31.25 | 2 | 2 | 31.25 | 22.57 | NR | U | N/A
|  | Unknown | 33.104 | 18 | 2 | 2 | 18 | 9.35 | NR | U | N/A
| Syringe S1 | Benzyl Alcohol | 13.479 | 3152.28 | 2 | 2 | 3152.28 | 1158.8 | NR | IC | 100-51-6
|  | C14 Alkene | 17.651 | 11.99 | 2 | 2 | 11.99 | -22.74 | NR | U | N/A
|  | C14 Alkene | 17.788 | 46.82 | 2 | 2 | 46.82 | -78.09 | NR | U | N/A
|  | C14 Alkene | 17.868 | 95.7 | 2 | 2 | 95.7 | -60.34 | NR | U | N/A
|  | C14 Alkene | 18.165 | 253.11 | 2 | 2 | 253.11 | -122.54 | NR | U | N/A
|  | C14 Alkane | 18.245 | 195.69 | 2 | 2 | 195.69 | -104.68 | NR | U | N/A
|  | C14 Alkane | 18.337 | 151.64 | 2 | 2 | 151.64 | -98.69 | NR | U | N/A
|  | Alkene | 18.76 | 17.07 | 2 | 2 | 17.07 | 4.61 | NR | U | N/A
|  | Alkene | 18.92 | 50.55 | 2 | 2 | 50.55 | 35.93 | NR | U | N/A
|  | 7-ethyl-2-methyl-4-Undecanol | 21.423 | 745.1 | 2 | 2 | 745.1 | 179.39 | NR | MPC | 103-20-8
|  | 9,10-dihydro-Anthracene | 24.829 | 5.37 | 2 | 2 | 5.37 | -23.56 | NR | MPC | 613-31-0
|  | Unknown | 26.372 | 1.98 | 2 | 2 | 1.98 | -5.91 | NR | U | N/A
|  | C16-C20 Alkyl Benzene | 27.778 | 69.24 | 2 | 2 | 69.24 | -1.25 | NR | U | N/A
|  | Unknown | 29.104 | 34.27 | 2 | 2 | 34.27 | 29.9 | NR | U | N/A
|  | Silane | 30.658 | 29.83 | 2 | 2 | 29.83 | 23.33 | NR | U | N/A
|  | 2-(phenylmethyl)-Benzenemethanol | 31.184 | 12.12 | 2 | 2 | 12.12 | 4.09 | NR | MPC | 1586-00-1
|  | Bibenzyl | 31.584 | 5.62 | 2 | 2 | 5.62 | -1.63 | NR | TIC | 103-29-7
|  | C20-C25 Phthalate | 31.767 | 41.79 | 2 | 2 | 41.79 | 33.11 | NR | U | N/A
|  | Unknown | 32.121 | 7.82 | 2 | 2 | 7.82 | 7.82 | NR | U | N/A
|  | Unknown | 33.092 | 30.04 | 2 | 2 | 30.04 | 23.76 | NR | U | N/A
| Syringe S2 | Benzyl Alcohol | 13.491 | 2760.3 | 2 | 2 | 2760.3 | 766.82 | NR | IC | 100-51-6
|  | C14 Alkene | 17.662 | 11.39 | 2 | 2 | 11.39 | -23.34 | NR | U | N/A
|  | C14 Alkene | 17.788 | 42.88 | 2 | 2 | 42.88 | -82.03 | NR | U | N/A
|  | C14 Alkene | 17.868 | 79.68 | 2 | 2 | 79.68 | -76.36 | NR | U | N/A
|  | C14 Alkene | 18.165 | 209.24 | 2 | 2 | 209.24 | -166.41 | NR | U | N/A
|  | C14 Alkane | 18.245 | 161.25 | 2 | 2 | 161.25 | -139.12 | NR | U | N/A
|  | C14 Alkane | 18.348 | 128.62 | 2 | 2 | 128.62 | -121.71 | NR | U | N/A
|  | Alkene | 18.76 | 16.26 | 2 | 2 | 16.26 | 3.8 | NR | U | N/A
|  | Alkene | 18.92 | 50.72 | 2 | 2 | 50.72 | 36.1 | NR | U | N/A
|  | 7-ethyl-2-methyl-4-Undecanol | 21.411 | 795.86 | 2 | 2 | 795.86 | 230.15 | NR | MPC | 103-20-8
|  | 9,10-dihydro-Anthracene | 24.829 | 18.91 | 2 | 2 | 18.91 | -10.02 | NR | MPC | 613-31-0
|  | Unknown | 26.326 | 21.53 | 2 | 2 | 21.53 | 13.64 | NR | U | N/A
|  | C16-C20 Alkyl Benzene | 27.778 | 73.01 | 2 | 2 | 73.01 | 2.52 | NR | U | N/A
|  | Unknown | 29.092 | 20.65 | 2 | 2 | 20.65 | 16.28 | NR | U | N/A
|  | Silane | 30.658 | 32.78 | 2 | 2 | 32.78 | 26.28 | NR | U | N/A
|  | 2-(phenylmethyl)-Benzenemethanol | 31.184 | 22.37 | 2 | 2 | 22.37 | 14.34 | NR | MPC | 1586-00-1
|  | Bibenzyl | 31.584 | 10.32 | 2 | 2 | 10.32 | 3.07 | NR | TIC | 103-29-7
|  | C20-C25 Phthalate | 31.767 | 44.16 | 2 | 2 | 44.16 | 35.48 | NR | U | N/A
|  | Unknown | 32.121 | 16.6 | 2 | 2 | 16.6 | 16.6 | NR | U | N/A
|  | Unknown | 33.093 | 20.84 | 2 | 2 | 20.84 | 14.56 | NR | U | N/A
| Syringe S3 | Benzyl Alcohol | 13.491 | 2829.19 | 2 | 2 | 2829.19 | 835.71 | NR | IC | 100-51-6
|  | C14 Alkene | 17.651 | 10.92 | 2 | 2 | 10.92 | -23.81 | NR | U | N/A
|  | C14 Alkene | 17.788 | 42.63 | 2 | 2 | 42.63 | -82.28 | NR | U | N/A
|  | C14 Alkene | 17.868 | 89.78 | 2 | 2 | 89.78 | -66.26 | NR | U | N/A
|  | C14 Alkene | 18.166 | 238.25 | 2 | 2 | 238.25 | -137.4 | NR | U | N/A
|  | C14 Alkane | 18.246 | 183.71 | 2 | 2 | 183.71 | -116.66 | NR | U | N/A
|  | C14 Alkane | 18.348 | 144.08 | 2 | 2 | 144.08 | -106.25 | NR | U | N/A
|  | Alkene | 18.76 | 15.97 | 2 | 2 | 15.97 | 3.51 | NR | U | N/A
|  | Alkene | 18.92 | 48.65 | 2 | 2 | 48.65 | 34.03 | NR | U | N/A
|  | 7-ethyl-2-methyl-4-Undecanol | 21.412 | 822.43 | 2 | 2 | 822.43 | 256.72 | NR | MPC | 103-20-8
|  | 9,10-dihydro-Anthracene | 24.829 | 5.9 | 2 | 2 | 5.9 | -23.03 | NR | MPC | 613-31-0
|  | Unknown | 26.372 | 1.67 | 2 | 2 | 1.67 | -6.22 | NR | U | N/A
|  | C16-C20 Alkyl Benzene | 27.778 | 80.37 | 2 | 2 | 80.37 | 9.88 | NR | U | N/A
|  | Unknown | 29.092 | 4.24 | 2 | 2 | 4.24 | -0.13 | NR | U | N/A
|  | Silane | 30.658 | 35.17 | 2 | 2 | 35.17 | 28.67 | NR | U | N/A
|  | 2-(phenylmethyl)-Benzenemethanol | 31.184 | 0.36 | 2 | 2 | 0.36 | -7.67 | NR | MPC | 1586-00-1
|  | Bibenzyl | 31.595 | 0.03 | 2 | 2 | 0.03 | -7.22 | NR | TIC | 103-29-7
|  | C20-C25 Phthalate | 31.767 | 48.01 | 2 | 2 | 48.01 | 39.33 | NR | U | N/A
|  | Unknown | 32.098 | 0.31 | 2 | 2 | 0.31 | 0.31 | NR | U | N/A
|  | Unknown | 33.093 | 77.85 | 2 | 2 | 77.85 | 71.57 | NR | U | N/A
16-00866-N1 |  | SVOC
[8 empty] | Blank | Reportable
|  |  |  | Instrument | Concentration | Dilution | Corrected | Corrected | Not Reportable
|  | Compound | Ret Time | Result (ug/mL) | Factor | Factor | Result (ug/mL) | Result (ug/mL) | (R/NR) | ID | CAS #
1% S.A. | mb1603047OE | hexamethyl-Cyclotrisiloxane | 7.981 | 1.56 | 2 | 2 | 1.56 | 1.56 | NR | MPC | 541-05-9
30 minute |  | Benzyl Alcohol | 13.502 | 2887.53 | 2 | 2 | 2887.53 | 2887.53 | R | IC | 100-51-6
|  | decamethyl-Cyclopentasiloxane | 15.045 | 7.6 | 2 | 2 | 7.6 | 7.6 | NR | MPC | 541-02-6
|  | C14 Alkene | 17.651 | 1.69 | 2 | 2 | 1.69 | 1.69 | NR | U | N/A
|  | C14 Alkene | 17.8 | 1.66 | 2 | 2 | 1.66 | 1.66 | NR | U | N/A
|  | C14 Alkene | 17.88 | 9.8 | 2 | 2 | 9.8 | 9.8 | NR | U | N/A
|  | C14 Alkene | 18.165 | 29.04 | 2 | 2 | 29.04 | 29.04 | R | U | N/A
|  | C14 Alkane | 18.257 | 18.56 | 2 | 2 | 18.56 | 18.56 | R | U | N/A
|  | C14 Alkane | 18.348 | 13.09 | 2 | 2 | 13.09 | 13.09 | R | U | N/A
|  | Alkene | 18.942 | 1.97 | 2 | 2 | 1.97 | 1.97 | NR | U | N/A
|  | octamethyl-Cyclotetrasiloxane | 20.417 | 1.95 | 2 | 2 | 1.95 | 1.95 | NR | TIC | 556-67-2
|  | Diphenylmethane | 21.023 | 1.37 | 2 | 2 | 1.37 | 1.37 | NR | MPC | 101-85-5
|  | Unknown | 21.217 | 3.23 | 2 | 2 | 3.23 | 3.23 | NR | U | N/A
|  | 7-ethyl-2-methyl-4-Undecanol | 21.446 | 644.42 | 2 | 2 | 644.42 | 644.42 | R | MPC | 103-20-8
|  | 1-methyl-4-(phenylmethyl)-Benzene | 22.76 | 0.83 | 2 | 2 | 0.83 | 0.83 | NR | MPC | 620-83-7
|  | 9,10-dihydro-Anthracene | 24.84 | 3.63 | 2 | 2 | 3.63 | 3.63 | NR | MPC | 613-31-0
|  | cis-Stilbene | 25.686 | 10.9 | 2 | 2 | 10.9 | 10.9 | R | TIC | 645-49-8
|  | C16-C20 Alkyl Benzene | 27.789 | 3.19 | 2 | 2 | 3.19 | 3.19 | NR | U | N/A
|  | Silane | 30.692 | 9.26 | 2 | 2 | 9.26 | 9.26 | NR | U | N/A
|  | C20-C25 Phthalate | 31.778 | 77.09 | 2 | 2 | 77.09 | 77.09 | R | U | N/A
|  | Unknown | 33.104 | 67.2 | 2 | 2 | 67.2 | 67.2 | R | U | N/A
|  | Unknown | 34.693 | 18.7 | 2 | 2 | 18.7 | 18.7 | R | U | N/A
| Control | hexamethyl-Cyclotrisiloxane | 7.936 | 23.14 | 2 | 2 | 23.14 | 21.58 | NR | MPC | 541-05-9
|  | Benzyl Alcohol | 13.491 | 2807.33 | 2 | 2 | 2807.33 | -80.2 | NR | IC | 100-51-6
|  | decamethyl-Cyclopentasiloxane | 15.034 | 22.12 | 2 | 2 | 22.12 | 14.52 | NR | MPC | 541-02-6
|  | C14 Alkene | 17.674 | 17.81 | 2 | 2 | 17.81 | 16.12 | NR | U | N/A
|  | C14 Alkene | 17.8 | 72.34 | 2 | 2 | 72.34 | 70.68 | NR | U | N/A
|  | C14 Alkene | 17.88 | 102.89 | 2 | 2 | 102.89 | 93.09 | NR | U | N/A
|  | C14 Alkene | 18.177 | 260.44 | 2 | 2 | 260.44 | 231.4 | NR | U | N/A
|  | C14 Alkane | 18.268 | 213.32 | 2 | 2 | 213.32 | 194.76 | NR | U | N/A
|  | C14 Alkane | 18.36 | 181.21 | 2 | 2 | 181.21 | 168.12 | NR | U | N/A
|  | Cycloalkane | 18.508 | 8.74 | 2 | 2 | 8.74 | 8.74 | NR | U | N/A
|  | Cycloalkane | 18.588 | 9.98 | 2 | 2 | 9.98 | 9.98 | NR | U | N/A
|  | Branched C12 Alkene | 18.68 | 10.91 | 2 | 2 | 10.91 | 10.91 | R | U | N/A
|  | Branched C12 Alkene | 18.725 | 8.08 | 2 | 2 | 8.08 | 8.08 | NR | U | N/A
|  | Alkene | 18.782 | 24.85 | 2 | 2 | 24.85 | 24.85 | R | U | N/A
|  | Alkene | 18.943 | 93.85 | 2 | 2 | 93.85 | 91.88 | NR | U | N/A
|  | octamethyl-Cyclotetrasiloxane | 20.383 | 19.58 | 2 | 2 | 19.58 | 17.63 | NR | TIC | 556-67-2
|  | Diphenylmethane | 21.023 | 7.01 | 2 | 2 | 7.01 | 5.64 | NR | MPC | 101-85-5
|  | Unknown | 21.217 | 7.44 | 2 | 2 | 7.44 | 7.44 | NR | U | N/A
|  | 7-ethyl-2-methyl-4-Undecanol | 21.423 | 596.25 | 2 | 2 | 596.25 | -48.17 | NR | MPC | 103-20-8
|  | Siloxane | 22.554 | 7.64 | 2 | 2 | 7.64 | 7.64 | NR | U | N/A
|  | 9,10-dihydro-Anthracene | 24.84 | 19.23 | 2 | 2 | 19.23 | 15.6 | NR | MPC | 613-31-0
|  | cis-Stilbene | 25.686 | 8.42 | 2 | 2 | 8.42 | 8.42 | NR | TIC | 645-49-8
|  | C16-C20 Alkyl Benzene | 27.789 | 45.16 | 2 | 2 | 45.16 | 41.97 | NR | U | N/A
|  | Silane | 30.669 | 35.05 | 2 | 2 | 35.05 | 25.79 | NR | U | N/A
|  | C20-C25 Phthalate | 31.778 | 177.08 | 2 | 2 | 177.08 | 99.99 | NR | U | N/A
|  | Unknown | 33.104 | 224.71 | 2 | 2 | 224.71 | 157.51 | NR | U | N/A
|  | Unknown | 34.693 | 124.04 | 2 | 2 | 124.04 | 105.34 | NR | U | N/A
|  | Phthalate | 38.51 | 5.27 | 2 | 2 | 5.27 | 5.27 | NR | U | N/A
| Catheter/Syringe S1 | hexamethyl-Cyclotrisiloxane | 7.936 | 18.11 | 2 | 2 | 18.11 | 16.55 | NR | MPC | 541-05-9
|  | Benzyl Alcohol | 13.502 | 3127.22 | 2 | 2 | 3127.22 | 239.69 | NR | IC | 100-51-6
|  | decamethyl-Cyclopentasiloxane | 15.034 | 23.37 | 2 | 2 | 23.37 | 15.77 | NR | MPC | 541-02-6
|  | C14 Alkene | 17.8 | 34.58 | 2 | 2 | 34.58 | 32.92 | NR | U | N/A
|  | C14 Alkene | 17.88 | 62.94 | 2 | 2 | 62.94 | 53.14 | NR | U | N/A
|  | C14 Alkene | 18.177 | 182.49 | 2 | 2 | 182.49 | 153.45 | NR | U | N/A
|  | C14 Alkane | 18.268 | 136.09 | 2 | 2 | 136.09 | 117.53 | NR | U | N/A
|  | C14 Alkane | 18.36 | 142.92 | 2 | 2 | 142.92 | 129.83 | NR | U | N/A
|  | Alkene | 18.783 | 11.11 | 2 | 2 | 11.11 | 11.11 | NR | U | N/A
|  | Alkene | 18.943 | 52.62 | 2 | 2 | 52.62 | 50.65 | NR | U | N/A
|  | octamethyl-Cyclotetrasiloxane | 20.394 | 31.93 | 2 | 2 | 31.93 | 29.98 | NR | TIC | 556-67-2
|  | 7-ethyl-2-methyl-4-Undecanol | 21.434 | 628.5 | 2 | 2 | 628.5 | -15.92 | NR | MPC | 103-20-8
|  | cis-Stilbene | 25.686 | 10.66 | 2 | 2 | 10.66 | 10.66 | NR | TIC | 645-49-8
|  | C16-C20 Alkyl Benzene | 27.789 | 16.16 | 2 | 2 | 16.16 | 12.97 | NR | U | N/A
|  | Silane | 30.669 | 27.86 | 2 | 2 | 27.86 | 18.6 | NR | U | N/A
|  | C20-C25 Phthalate | 31.778 | 126.53 | 2 | 2 | 126.53 | 49.44 | NR | U | N/A
|  | Unknown | 33.104 | 241.71 | 2 | 2 | 241.71 | 174.51 | NR | U | N/A
|  | Unknown | 34.693 | 106.02 | 2 | 2 | 106.02 | 87.32 | NR | U | N/A
| Catheter/Syringe S2 | hexamethyl-Cyclotrisiloxane | 7.936 | 5.97 | 2 | 2 | 5.97 | 4.41 | NR | MPC | 541-05-9
|  | Benzyl Alcohol | 13.502 | 3183.49 | 2 | 2 | 3183.49 | 295.96 | NR | IC | 100-51-6
|  | decamethyl-Cyclopentasiloxane | 15.034 | 8.89 | 2 | 2 | 8.89 | 1.29 | NR | MPC | 541-02-6
|  | C14 Alkene | 17.799 | 12.03 | 2 | 2 | 12.03 | 10.37 | NR | U | N/A
|  | C14 Alkene | 17.88 | 34.03 | 2 | 2 | 34.03 | 24.23 | NR | U | N/A
|  | C14 Alkene | 18.165 | 96.31 | 2 | 2 | 96.31 | 67.27 | NR | U | N/A
|  | C14 Alkane | 18.257 | 68.49 | 2 | 2 | 68.49 | 49.93 | NR | U | N/A
|  | C14 Alkane | 18.348 | 53.7 | 2 | 2 | 53.7 | 40.61 | NR | U | N/A
|  | Alkene | 18.782 | 3.97 | 2 | 2 | 3.97 | 3.97 | NR | U | N/A
|  | Alkene | 18.942 | 17.37 | 2 | 2 | 17.37 | 15.4 | NR | U | N/A
|  | octamethyl-Cyclotetrasiloxane | 20.44 | 7.99 | 2 | 2 | 7.99 | 6.04 | NR | TIC | 556-67-2
|  | 7-ethyl-2-methyl-4-Undecanol | 21.446 | 652.79 | 2 | 2 | 652.79 | 8.37 | NR | MPC | 103-20-8
|  | cis-Stilbene | 25.686 | 8.57 | 2 | 2 | 8.57 | 8.57 | NR | TIC | 645-49-8
|  | C16-C20 Alkyl Benzene | 27.789 | 13.63 | 2 | 2 | 13.63 | 10.44 | NR | U | N/A
|  | Silane | 30.669 | 16.19 | 2 | 2 | 16.19 | 6.93 | NR | U | N/A
|  | C20-C25 Phthalate | 31.778 | 52.66 | 2 | 2 | 52.66 | -24.43 | NR | U | N/A
|  | Unknown | 33.104 | 135.94 | 2 | 2 | 135.94 | 68.74 | NR | U | N/A
|  | Unknown | 34.693 | 31.45 | 2 | 2 | 31.45 | 12.75 | NR | U | N/A
| Catheter/Syringe S3 | hexamethyl-Cyclotrisiloxane | 7.936 | 10.42 | 2 | 2 | 10.42 | 8.86 | NR | MPC | 541-05-9
|  | Benzyl Alcohol | 13.502 | 3418.34 | 2 | 2 | 3418.34 | 530.81 | NR | IC | 100-51-6
|  | decamethyl-Cyclopentasiloxane | 15.034 | 13.84 | 2 | 2 | 13.84 | 6.24 | NR | MPC | 541-02-6
|  | C14 Alkene | 17.8 | 23.96 | 2 | 2 | 23.96 | 22.3 | NR | U | N/A
|  | C14 Alkene | 17.88 | 54.36 | 2 | 2 | 54.36 | 44.56 | NR | U | N/A
|  | C14 Alkene | 18.165 | 152.7 | 2 | 2 | 152.7 | 123.66 | NR | U | N/A
|  | C14 Alkane | 18.257 | 112.56 | 2 | 2 | 112.56 | 94 | NR | U | N/A
|  | C14 Alkane | 18.36 | 90.02 | 2 | 2 | 90.02 | 76.93 | NR | U | N/A
|  | Alkene | 18.783 | 7.78 | 2 | 2 | 7.78 | 7.78 | NR | U | N/A
|  | Alkene | 18.943 | 34.51 | 2 | 2 | 34.51 | 32.54 | NR | U | N/A
|  | octamethyl-Cyclotetrasiloxane | 20.429 | 13.53 | 2 | 2 | 13.53 | 11.58 | NR | TIC | 556-67-2
|  | 7-ethyl-2-methyl-4-Undecanol | 21.446 | 679.18 | 2 | 2 | 679.18 | 34.76 | NR | MPC | 103-20-8
|  | cis-Stilbene | 25.686 | 7.95 | 2 | 2 | 7.95 | 7.95 | NR | TIC | 645-49-8
|  | C16-C20 Alkyl Benzene | 27.789 | 17.96 | 2 | 2 | 17.96 | 14.77 | NR | U | N/A
|  | Silane | 30.669 | 20.53 | 2 | 2 | 20.53 | 11.27 | NR | U | N/A
|  | C20-C25 Phthalate | 31.778 | 70.81 | 2 | 2 | 70.81 | -6.28 | NR | U | N/A
|  | Unknown | 33.104 | 171.7 | 2 | 2 | 171.7 | 104.5 | NR | U | N/A
|  | Unknown | 34.681 | 43.69 | 2 | 2 | 43.69 | 24.99 | NR | U | N/A
| Syringe S1 | hexamethyl-Cyclotrisiloxane | 7.936 | 51.81 | 2 | 2 | 51.81 | 50.25 | NR | MPC | 541-05-9
|  | Benzyl Alcohol | 13.456 | 2755.56 | 2 | 2 | 2755.56 | -131.97 | NR | IC | 100-51-6
|  | decamethyl-Cyclopentasiloxane | 15.022 | 48.4 | 2 | 2 | 48.4 | 40.8 | NR | MPC | 541-02-6
|  | C14 Alkene | 17.674 | 33.1 | 2 | 2 | 33.1 | 31.41 | NR | U | N/A
|  | C14 Alkene | 17.8 | 129.66 | 2 | 2 | 129.66 | 128 | NR | U | N/A
|  | C14 Alkene | 17.88 | 161.54 | 2 | 2 | 161.54 | 151.74 | NR | U | N/A
|  | C14 Alkene | 18.177 | 384.19 | 2 | 2 | 384.19 | 355.15 | NR | U | N/A
|  | C14 Alkane | 18.268 | 352.14 | 2 | 2 | 352.14 | 333.58 | NR | U | N/A
|  | C14 Alkane | 18.36 | 309.59 | 2 | 2 | 309.59 | 296.5 | NR | U | N/A
|  | Alkene | 18.508 | 15.73 | 2 | 2 | 15.73 | 15.73 | NR | U | N/A
|  | Alkene | 18.588 | 15.99 | 2 | 2 | 15.99 | 14.02 | NR | U | N/A
|  | Branched C12 Alkene | 18.68 | 17.21 | 2 | 2 | 17.21 | 17.21 | NR | U | N/A
|  | Branched C12 Alkene | 18.725 | 12 | 2 | 2 | 12 | 12 | NR | U | N/A
|  | Alkene | 18.783 | 35.82 | 2 | 2 | 35.82 | 35.82 | NR | U | N/A
|  | Alkene | 18.943 | 170.38 | 2 | 2 | 170.38 | 168.41 | NR | U | N/A
|  | octamethyl-Cyclotetrasiloxane | 20.383 | 107.82 | 2 | 2 | 107.82 | 105.87 | NR | TIC | 556-67-2
|  | Diphenylmethane | 21.023 | 15.31 | 2 | 2 | 15.31 | 13.94 | NR | MPC | 101-85-5
|  | Unknown | 21.206 | 10.64 | 2 | 2 | 10.64 | 10.64 | NR | U | N/A
|  | 7-ethyl-2-methyl-4-Undecanol | 21.4 | 678.53 | 2 | 2 | 678.53 | 34.11 | NR | MPC | 103-20-8
|  | Siloxane | 22.589 | 22.56 | 2 | 2 | 22.56 | 22.56 | NR | U | N/A
|  | 9,10-dihydro-Anthracene | 24.84 | 37.69 | 2 | 2 | 37.69 | 34.06 | NR | MPC | 613-31-0
|  | C16-C20 Alkyl Benzene | 27.789 | 17.02 | 2 | 2 | 17.02 | 13.83 | NR | U | N/A
|  | Silane | 30.669 | 35.68 | 2 | 2 | 35.68 | 26.42 | NR | U | N/A
|  | C20-C25 Phthalate | 31.778 | 92.07 | 2 | 2 | 92.07 | 14.98 | NR | U | N/A
|  | Unknown | 33.104 | 305.26 | 2 | 2 | 305.26 | 238.06 | NR | U | N/A
|  | Unknown | 34.693 | 52.28 | 2 | 2 | 52.28 | 33.58 | NR | U | N/A
|  | Phthalate | 38.51 | 7.1 | 2 | 2 | 7.1 | 7.1 | NR | U | N/A
| Syringe S2 | hexamethyl-Cyclotrisiloxane | 7.936 | 36.43 | 2 | 2 | 36.43 | 34.87 | NR | MPC | 541-05-9
|  | Benzyl Alcohol | 13.502 | 3510.61 | 2 | 2 | 3510.61 | 623.08 | NR | IC | 100-51-6
|  | decamethyl-Cyclopentasiloxane | 15.034 | 22.24 | 2 | 2 | 22.24 | 14.64 | NR | MPC | 541-02-6
|  | C14 Alkene | 17.674 | 11.9 | 2 | 2 | 11.9 | 10.21 | NR | U | N/A
|  | C14 Alkene | 17.799 | 49.78 | 2 | 2 | 49.78 | 48.12 | NR | U | N/A
|  | C14 Alkene | 17.88 | 78.62 | 2 | 2 | 78.62 | 68.82 | NR | U | N/A
|  | C14 Alkene | 18.177 | 219.59 | 2 | 2 | 219.59 | 190.55 | NR | U | N/A
|  | C14 Alkane | 18.268 | 171.1 | 2 | 2 | 171.1 | 152.54 | NR | U | N/A
|  | C14 Alkane | 18.36 | 162.9 | 2 | 2 | 162.9 | 149.81 | NR | U | N/A
|  | Alkene | 18.508 | 10.76 | 2 | 2 | 10.76 | 10.76 | NR | U | N/A
|  | Alkene | 18.588 | 6.67 | 2 | 2 | 6.67 | 4.7 | NR | U | N/A
|  | Branched C12 Alkene | 18.68 | 5.82 | 2 | 2 | 5.82 | 5.82 | NR | U | N/A
|  | Branched C12 Alkene | 18.725 | 4.02 | 2 | 2 | 4.02 | 4.02 | NR | U | N/A
|  | Alkene | 18.782 | 13.52 | 2 | 2 | 13.52 | 13.52 | NR | U | N/A
|  | Alkene | 18.942 | 68.7 | 2 | 2 | 68.7 | 66.73 | NR | U | N/A
|  | octamethyl-Cyclotetrasiloxane | 20.417 | 38.43 | 2 | 2 | 38.43 | 36.48 | NR | TIC | 556-67-2
|  | Diphenylmethane | 21.023 | 2.82 | 2 | 2 | 2.82 | 1.45 | NR | MPC | 101-85-5
|  | Unknown | 21.206 | 8.02 | 2 | 2 | 8.02 | 8.02 | NR | U | N/A
|  | 7-ethyl-2-methyl-4-Undecanol | 21.423 | 694.46 | 2 | 2 | 694.46 | 50.04 | NR | MPC | 103-20-8
|  | Siloxane | 22.589 | 5.19 | 2 | 2 | 5.19 | 5.19 | NR | U | N/A
|  | 9,10-dihydro-Anthracene | 24.84 | 6.4 | 2 | 2 | 6.4 | 2.77 | NR | MPC | 613-31-0
|  | C16-C20 Alkyl Benzene | 27.789 | 27.56 | 2 | 2 | 27.56 | 24.37 | NR | U | N/A
|  | Silane | 30.669 | 56.63 | 2 | 2 | 56.63 | 47.37 | NR | U | N/A
|  | C20-C25 Phthalate | 31.778 | 94.87 | 2 | 2 | 94.87 | 17.78 | NR | U | N/A
|  | Unknown | 33.104 | 309.4 | 2 | 2 | 309.4 | 242.2 | NR | U | N/A
|  | Unknown | 34.693 | 40.66 | 2 | 2 | 40.66 | 21.96 | NR | U | N/A
|  | Phthalate | 38.51 | 10.26 | 2 | 2 | 10.26 | 10.26 | NR | U | N/A
| Syringe S3 | hexamethyl-Cyclotrisiloxane | 7.936 | 38.23 | 2 | 2 | 38.23 | 36.67 | NR | MPC | 541-05-9
|  | Benzyl Alcohol | 13.513 | 3515.39 | 2 | 2 | 3515.39 | 627.86 | NR | IC | 100-51-6
|  | decamethyl-Cyclopentasiloxane | 15.034 | 19.67 | 2 | 2 | 19.67 | 12.07 | NR | MPC | 541-02-6
|  | C14 Alkene | 17.674 | 12.93 | 2 | 2 | 12.93 | 11.24 | NR | U | N/A
|  | C14 Alkene | 17.799 | 48.73 | 2 | 2 | 48.73 | 47.07 | NR | U | N/A
|  | C14 Alkene | 17.88 | 81.69 | 2 | 2 | 81.69 | 71.89 | NR | U | N/A
|  | C14 Alkene | 18.177 | 219.76 | 2 | 2 | 219.76 | 190.72 | NR | U | N/A
|  | C14 Alkane | 18.268 | 166.9 | 2 | 2 | 166.9 | 148.34 | NR | U | N/A
|  | C14 Alkane | 18.36 | 143.22 | 2 | 2 | 143.22 | 130.13 | NR | U | N/A
|  | Alkene | 18.508 | 1.83 | 2 | 2 | 1.83 | 1.83 | NR | U | N/A
|  | Alkene | 18.588 | 4.49 | 2 | 2 | 4.49 | 2.52 | NR | U | N/A
|  | Branched C12 Alkene | 18.68 | 3.91 | 2 | 2 | 3.91 | 3.91 | R | U | N/A
|  | Branched C12 Alkene | 18.725 | 3.02 | 2 | 2 | 3.02 | 3.02 | NR | U | N/A
|  | Alkene | 18.782 | 13.91 | 2 | 2 | 13.91 | 13.91 | R | U | N/A
|  | Alkene | 18.942 | 64.61 | 2 | 2 | 64.61 | 62.64 | NR | U | N/A
|  | octamethyl-Cyclotetrasiloxane | 20.428 | 21.88 | 2 | 2 | 21.88 | 19.93 | NR | TIC | 556-67-2
|  | Diphenylmethane | 21.023 | 3.92 | 2 | 2 | 3.92 | 2.55 | NR | MPC | 101-85-5
|  | Unknown | 21.206 | 7.25 | 2 | 2 | 7.25 | 7.25 | NR | U | N/A
|  | 7-ethyl-2-methyl-4-Undecanol | 21.423 | 602.24 | 2 | 2 | 602.24 | -42.18 | NR | MPC | 103-20-8
|  | Siloxane | 22.76 | 1.69 | 2 | 2 | 1.69 | 1.69 | NR | U | N/A
|  | 9,10-dihydro-Anthracene | 24.84 | 9.99 | 2 | 2 | 9.99 | 6.36 | NR | MPC | 613-31-0
|  | C16-C20 Alkyl Benzene | 27.789 | 38.27 | 2 | 2 | 38.27 | 35.08 | NR | U | N/A
|  | Silane | 30.669 | 71.43 | 2 | 2 | 71.43 | 62.17 | NR | U | N/A
|  | C20-C25 Phthalate | 31.778 | 104.78 | 2 | 2 | 104.78 | 27.69 | NR | U | N/A
|  | Unknown | 33.104 | 267.21 | 2 | 2 | 267.21 | 200.01 | NR | U | N/A
|  | Unknown | 34.693 | 27.25 | 2 | 2 | 27.25 | 8.55 | NR | U | N/A
|  | Phthalate | 38.51 | 15.39 | 2 | 2 | 15.39 | 15.39 | NR | U | N/A
16-00866-N1 |  | SVOC
[8 empty] | Blank | Reportable
|  |  |  | Instrument | Concentration | Dilution | Corrected | Corrected | Not Reportable
|  | Compound | Ret Time | Result (ug/mL) | Factor | Factor | Result (ug/mL) | Result (ug/mL) | (R/NR) | ID | CAS #
3% S.A. | mb1603053OE | Benzyl Alcohol | 13.376 | 2530.17 | 1 | 2 | 5060.34 | 5060.34 | R | IC | 100-51-6
10 minute |  | decamethyl-Cyclopentasiloxane | 15.011 | 23.35 | 1 | 2 | 46.7 | 46.7 | R | MPC | 541-02-6
|  | C14 Alkene | 17.64 | 14.77 | 1 | 2 | 29.54 | 29.54 | R | U | N/A
|  | C14 Alkene | 17.765 | 62.08 | 1 | 2 | 124.16 | 124.16 | R | U | N/A
|  | C14 Alkene | 17.845 | 143.22 | 1 | 2 | 286.44 | 286.44 | R | U | N/A
|  | C14 Alkene | 18.143 | 352.72 | 1 | 2 | 705.44 | 705.44 | R | U | N/A
|  | C14 Alkane | 18.234 | 266.88 | 1 | 2 | 533.76 | 533.76 | R | U | N/A
|  | C14 Alkane | 18.325 | 216.73 | 1 | 2 | 433.46 | 433.46 | R | U | N/A
|  | Alkene | 18.474 | 4.51 | 1 | 2 | 9.02 | 9.02 | NR | U | N/A
|  | Alkene | 18.554 | 4.88 | 1 | 2 | 9.76 | 9.76 | NR | U | N/A
|  | Branched C12 Alkene | 18.645 | 5.62 | 1 | 2 | 11.24 | 11.24 | NR | U | N/A
|  | Alkene | 18.748 | 24.67 | 1 | 2 | 49.34 | 49.34 | R | U | N/A
|  | Alkene | 18.908 | 70.38 | 1 | 2 | 140.76 | 140.76 | R | U | N/A
|  | 7-ethyl-2-methyl-4-Undecanol | 21.411 | 809.05 | 1 | 2 | 1618.1 | 1618.1 | R | MPC | 103-20-8
|  | C8 Alkyl Benzene | 21.491 | 73.51 | 1 | 2 | 147.02 | 147.02 | R | U | N/A
|  | 1,2-Dodecanediol | 21.777 | 10.12 | 1 | 2 | 20.24 | 20.24 | R | TIC | 1119-87-5
|  | cis-Stilbene | 25.652 | 11.78 | 1 | 2 | 23.56 | 23.56 | R | TIC | 645-49-8
|  | C16-C20 Alkyl Benzene | 27.766 | 53.02 | 1 | 2 | 106.04 | 106.04 | R | U | N/A
|  | Silane | 30.647 | 24.51 | 1 | 2 | 49.02 | 49.02 | R | U | N/A
|  | C20-C25 Phthalate | 31.755 | 53.6 | 1 | 2 | 107.2 | 107.2 | R | U | N/A
|  | Unknown | 33.07 | 110.27 | 1 | 2 | 220.54 | 220.54 | R | U | N/A
|  | Unknown | 34.647 | 27.29 | 1 | 2 | 54.58 | 54.58 | R | U | N/A
| Control | Benzyl Alcohol | 13.376 | 2433.79 | 1 | 2 | 4867.58 | -192.76 | NR | IC | 100-51-6
|  | decamethyl-Cyclopentasiloxane | 14.999 | 16.73 | 1 | 2 | 33.46 | -13.24 | NR | MPC | 541-02-6
|  | C14 Alkene | 17.64 | 31.93 | 1 | 2 | 63.86 | 34.32 | NR | U | N/A
|  | C14 Alkene | 17.765 | 124.6 | 1 | 2 | 249.2 | 125.04 | NR | U | N/A
|  | C14 Alkene | 17.857 | 210.66 | 1 | 2 | 421.32 | 134.88 | NR | U | N/A
|  | C14 Alkene | 17.994 | 13.59 | 1 | 2 | 27.18 | 27.18 | R | U | N/A
|  | C14 Alkene | 18.154 | 474.42 | 1 | 2 | 948.84 | 243.4 | NR | U | N/A
|  | C14 Alkane | 18.245 | 371.11 | 1 | 2 | 742.22 | 208.46 | NR | U | N/A
|  | C14 Alkane | 18.337 | 326 | 1 | 2 | 652 | 218.54 | NR | U | N/A
|  | Alkene | 18.417 | 22.78 | 1 | 2 | 45.56 | 45.56 | R | U | N/A
|  | Alkene | 18.474 | 20.69 | 1 | 2 | 41.38 | 32.36 | NR | U | N/A
|  | Alkene | 18.554 | 14.53 | 1 | 2 | 29.06 | 19.3 | NR | U | N/A
|  | Branched C12 Alkene | 18.634 | 14.75 | 1 | 2 | 29.5 | 18.26 | NR | U | N/A
|  | Alkene | 18.691 | 11.24 | 1 | 2 | 22.48 | 22.48 | R | U | N/A
|  | Alkene | 18.748 | 46.44 | 1 | 2 | 92.88 | 43.54 | NR | U | N/A
|  | Alkene | 18.908 | 136.09 | 1 | 2 | 272.18 | 131.42 | NR | U | N/A
|  | Alkene | 19.034 | 10.38 | 1 | 2 | 20.76 | 20.76 | R | U | N/A
|  | octamethyl-Cyclotetrasiloxane | 20.429 | 10.85 | 1 | 2 | 21.7 | 21.7 | R | TIC | 556-67-2
|  | 7-ethyl-2-methyl-4-Undecanol | 21.4 | 828.2 | 1 | 2 | 1656.4 | 38.3 | NR | MPC | 103-20-8
|  | cis-Stilbene | 25.652 | 8.2 | 1 | 2 | 16.4 | -7.16 | NR | TIC | 645-49-8
|  | C16-C20 Alkyl Benzene | 27.766 | 84.43 | 1 | 2 | 168.86 | 62.82 | NR | U | N/A
|  | Silane | 30.635 | 68.58 | 1 | 2 | 137.16 | 88.14 | NR | U | N/A
|  | C20-C25 Phthalate | 31.755 | 118.79 | 1 | 2 | 237.58 | 130.38 | NR | U | N/A
|  | Unknown | 33.07 | 204.12 | 1 | 2 | 408.24 | 187.7 | NR | U | N/A
|  | Unknown | 34.647 | 35.69 | 1 | 2 | 71.38 | 16.8 | NR | U | N/A
|  | Phthalate | 38.43 | 12.8 | 1 | 2 | 25.6 | 25.6 | R | U | N/A
| Catheter/Syringe S1 | Benzyl Alcohol | 13.376 | 2640.19 | 1 | 2 | 5280.38 | 220.04 | NR | IC | 100-51-6
|  | decamethyl-Cyclopentasiloxane | 15.011 | 11.25 | 1 | 2 | 22.5 | -24.2 | NR | MPC | 541-02-6
|  | C14 Alkene | 17.64 | 23.68 | 1 | 2 | 47.36 | 17.82 | NR | U | N/A
|  | C14 Alkene | 17.765 | 95.67 | 1 | 2 | 191.34 | 67.18 | NR | U | N/A
|  | C14 Alkene | 17.857 | 186.95 | 1 | 2 | 373.9 | 87.46 | NR | U | N/A
|  | C14 Alkene | 18.154 | 428.18 | 1 | 2 | 856.36 | 150.92 | NR | U | N/A
|  | C14 Alkane | 18.245 | 331.39 | 1 | 2 | 662.78 | 129.02 | NR | U | N/A
|  | C14 Alkane | 18.337 | 278.91 | 1 | 2 | 557.82 | 124.36 | NR | U | N/A
|  | Alkene | 18.474 | 10.43 | 1 | 2 | 20.86 | 11.84 | NR | U | N/A
|  | Alkene | 18.554 | 11.73 | 1 | 2 | 23.46 | 13.7 | NR | U | N/A
|  | Branched C12 Alkene | 18.645 | 12.04 | 1 | 2 | 24.08 | 12.84 | NR | U | N/A
|  | Alkene | 18.748 | 37.19 | 1 | 2 | 74.38 | 25.04 | NR | U | N/A
|  | Alkene | 18.908 | 110.77 | 1 | 2 | 221.54 | 80.78 | NR | U | N/A
|  | octamethyl-Cyclotetrasiloxane | 20.451 | 13.36 | 1 | 2 | 26.72 | 26.72 | NR | TIC | 556-67-2
|  | 7-ethyl-2-methyl-4-Undecanol | 21.412 | 974.4 | 1 | 2 | 1948.8 | 330.7 | NR | MPC | 103-20-8
|  | cis-Stilbene | 25.652 | 10.75 | 1 | 2 | 21.5 | -2.06 | NR | TIC | 645-49-8
|  | C16-C20 Alkyl Benzene | 27.766 | 31.45 | 1 | 2 | 62.9 | -43.14 | NR | U | N/A
|  | Silane | 30.635 | 67.55 | 1 | 2 | 135.1 | 86.08 | NR | U | N/A
|  | C20-C25 Phthalate | 31.755 | 113.76 | 1 | 2 | 227.52 | 120.32 | NR | U | N/A
|  | Unknown | 33.07 | 202.11 | 1 | 2 | 404.22 | 183.68 | NR | U | N/A
|  | Unknown | 34.647 | 31.68 | 1 | 2 | 63.36 | 8.78 | NR | U | N/A
|  | Phthalate | 38.43 | 12.62 | 1 | 2 | 25.24 | 25.24 | NR | U | N/A
| Catheter/Syringe S2 | Benzyl Alcohol | 13.388 | 2708.81 | 1 | 2 | 5417.62 | 357.28 | NR | IC | 100-51-6
|  | decamethyl-Cyclopentasiloxane | 15.011 | 8.96 | 1 | 2 | 17.92 | -28.78 | NR | MPC | 541-02-6
|  | C14 Alkene | 17.64 | 8.36 | 1 | 2 | 16.72 | -12.82 | NR | U | N/A
|  | C14 Alkene | 17.765 | 32.86 | 1 | 2 | 65.72 | -58.44 | NR | U | N/A
|  | C14 Alkene | 17.846 | 105.23 | 1 | 2 | 210.46 | -75.98 | NR | U | N/A
|  | C14 Alkene | 18.143 | 254.82 | 1 | 2 | 509.64 | -195.8 | NR | U | N/A
|  | C14 Alkane | 18.234 | 185.12 | 1 | 2 | 370.24 | -163.52 | NR | U | N/A
|  | C14 Alkane | 18.326 | 149.82 | 1 | 2 | 299.64 | -133.82 | NR | U | N/A
|  | Alkene | 18.463 | 2.28 | 1 | 2 | 4.56 | -4.46 | NR | U | N/A
|  | Alkene | 18.543 | 0.77 | 1 | 2 | 1.54 | -8.22 | NR | U | N/A
|  | Branched C12 Alkene | 18.634 | 0.5 | 1 | 2 | 1 | -10.24 | NR | U | N/A
|  | Alkene | 18.748 | 11.86 | 1 | 2 | 23.72 | -25.62 | NR | U | N/A
|  | Alkene | 18.908 | 33.52 | 1 | 2 | 67.04 | -73.72 | NR | U | N/A
|  | Alkene | 19.034 | 5.17 | 1 | 2 | 10.34 | 10.34 | NR | U | N/A
|  | octamethyl-Cyclotetrasiloxane | 20.429 | 2.13 | 1 | 2 | 4.26 | 4.26 | NR | TIC | 556-67-2
|  | 7-ethyl-2-methyl-4-Undecanol | 21.389 | 977.32 | 1 | 2 | 1954.64 | 336.54 | NR | MPC | 103-20-8
|  | cis-Stilbene | 25.652 | 5.7 | 1 | 2 | 11.4 | -12.16 | NR | TIC | 645-49-8
|  | C16-C20 Alkyl Benzene | 27.767 | 45.51 | 1 | 2 | 91.02 | -15.02 | NR | U | N/A
|  | Silane | 30.635 | 64.86 | 1 | 2 | 129.72 | 80.7 | NR | U | N/A
|  | C20-C25 Phthalate | 31.744 | 110.04 | 1 | 2 | 220.08 | 112.88 | NR | U | N/A
|  | Unknown | 33.07 | 138.53 | 1 | 2 | 277.06 | 56.52 | NR | U | N/A
|  | Unknown | 34.647 | 25.36 | 1 | 2 | 50.72 | -3.86 | NR | U | N/A
|  | Phthalate | 38.419 | 13.35 | 1 | 2 | 26.7 | 26.7 | NR | U | N/A
| Catheter/Syringe S3 | Benzyl Alcohol | 13.388 | 2851.43 | 1 | 2 | 5702.86 | 642.52 | NR | IC | 100-51-6
|  | decamethyl-Cyclopentasiloxane | 15.011 | 16.8 | 1 | 2 | 33.6 | -13.1 | NR | MPC | 541-02-6
|  | C14 Alkene | 17.64 | 13.63 | 1 | 2 | 27.26 | -2.28 | NR | U | N/A
|  | C14 Alkene | 17.765 | 50.07 | 1 | 2 | 100.14 | -24.02 | NR | U | N/A
|  | C14 Alkene | 17.845 | 155.81 | 1 | 2 | 311.62 | 25.18 | NR | U | N/A
|  | C14 Alkene | 18.142 | 343.65 | 1 | 2 | 687.3 | -18.14 | NR | U | N/A
|  | C14 Alkane | 18.234 | 267.15 | 1 | 2 | 534.3 | 0.54 | NR | U | N/A
|  | C14 Alkane | 18.325 | 231.43 | 1 | 2 | 462.86 | 29.4 | NR | U | N/A
|  | Alkene | 18.474 | 8.85 | 1 | 2 | 17.7 | 8.68 | NR | U | N/A
|  | Alkene | 18.554 | 7.38 | 1 | 2 | 14.76 | 5 | NR | U | N/A
|  | Branched C12 Alkene | 18.645 | 5.31 | 1 | 2 | 10.62 | -0.62 | NR | U | N/A
|  | Alkene | 18.748 | 26.22 | 1 | 2 | 52.44 | 3.1 | NR | U | N/A
|  | Alkene | 18.908 | 61 | 1 | 2 | 122 | -18.76 | NR | U | N/A
|  | octamethyl-Cyclotetrasiloxane | 20.463 | 11.73 | 1 | 2 | 23.46 | 23.46 | NR | TIC | 556-67-2
|  | 7-ethyl-2-methyl-4-Undecanol | 21.411 | 1016.7 | 1 | 2 | 2033.4 | 415.3 | NR | MPC | 103-20-8
|  | cis-Stilbene | 25.652 | 5.85 | 1 | 2 | 11.7 | -11.86 | NR | TIC | 645-49-8
|  | C16-C20 Alkyl Benzene | 27.766 | 47.28 | 1 | 2 | 94.56 | -11.48 | NR | U | N/A
|  | Silane | 30.635 | 79.74 | 1 | 2 | 159.48 | 110.46 | NR | U | N/A
|  | C20-C25 Phthalate | 31.755 | 123.21 | 1 | 2 | 246.42 | 139.22 | NR | U | N/A
|  | Unknown | 33.07 | 207.39 | 1 | 2 | 414.78 | 194.24 | NR | U | N/A
|  | Unknown | 34.647 | 33.58 | 1 | 2 | 67.16 | 12.58 | NR | U | N/A
|  | Phthalate | 38.419 | 13.48 | 1 | 2 | 26.96 | 26.96 | NR | U | N/A
| Syringe S1 | Benzyl Alcohol | 13.388 | 2907.36 | 1 | 2 | 5814.72 | 754.38 | NR | IC | 100-51-6
|  | decamethyl-Cyclopentasiloxane | 15.011 | 8.25 | 1 | 2 | 16.5 | -30.2 | NR | MPC | 541-02-6
|  | C14 Alkene | 17.64 | 19.75 | 1 | 2 | 39.5 | 9.96 | NR | U | N/A
|  | C14 Alkene | 17.765 | 77.97 | 1 | 2 | 155.94 | 31.78 | NR | U | N/A
|  | C14 Alkene | 17.845 | 184.14 | 1 | 2 | 368.28 | 81.84 | NR | U | N/A
|  | C14 Alkene | 18.143 | 418.42 | 1 | 2 | 836.84 | 131.4 | NR | U | N/A
|  | C14 Alkane | 18.234 | 318.24 | 1 | 2 | 636.48 | 102.72 | NR | U | N/A
|  | C14 Alkane | 18.325 | 264.24 | 1 | 2 | 528.48 | 95.02 | NR | U | N/A
|  | Alkene | 18.474 | 10.73 | 1 | 2 | 21.46 | 12.44 | NR | U | N/A
|  | Alkene | 18.554 | 10.18 | 1 | 2 | 20.36 | 10.6 | NR | U | N/A
|  | Branched C12 Alkene | 18.645 | 10.21 | 1 | 2 | 20.42 | 9.18 | NR | U | N/A
|  | Alkene | 18.748 | 31.59 | 1 | 2 | 63.18 | 13.84 | NR | U | N/A
|  | Alkene | 18.908 | 87.11 | 1 | 2 | 174.22 | 33.46 | NR | U | N/A
|  | octamethyl-Cyclotetrasiloxane | 20.463 | 17.25 | 1 | 2 | 34.5 | 34.5 | NR | TIC | 556-67-2
|  | 7-ethyl-2-methyl-4-Undecanol | 21.412 | 1159 | 1 | 2 | 2318 | 699.9 | NR | MPC | 103-20-8
|  | C16-C20 Alkyl Benzene | 27.755 | 47.11 | 1 | 2 | 94.22 | -11.82 | NR | U | N/A
|  | Silane | 30.635 | 95.18 | 1 | 2 | 190.36 | 141.34 | NR | U | N/A
|  | C20-C25 Phthalate | 31.755 | 139.31 | 1 | 2 | 278.62 | 171.42 | NR | U | N/A
|  | Unknown | 33.07 | 216.55 | 1 | 2 | 433.1 | 212.56 | NR | U | N/A
|  | Unknown | 34.647 | 32.7 | 1 | 2 | 65.4 | 10.82 | NR | U | N/A
|  | Phthalate | 38.419 | 15.24 | 1 | 2 | 30.48 | 30.48 | NR | U | N/A
| Syringe S2 | Benzyl Alcohol | 13.388 | 2795.64 | 1 | 2 | 5591.28 | 530.94 | NR | IC | 100-51-6
|  | decamethyl-Cyclopentasiloxane | 15.011 | 13.51 | 1 | 2 | 27.02 | -19.68 | NR | MPC | 541-02-6
|  | C14 Alkene | 17.64 | 22.99 | 1 | 2 | 45.98 | 16.44 | NR | U | N/A
|  | C14 Alkene | 17.765 | 88.75 | 1 | 2 | 177.5 | 53.34 | NR | U | N/A
|  | C14 Alkene | 17.845 | 190.86 | 1 | 2 | 381.72 | 95.28 | NR | U | N/A
|  | C14 Alkene | 18.154 | 419.79 | 1 | 2 | 839.58 | 134.14 | NR | U | N/A
|  | C14 Alkane | 18.246 | 328.17 | 1 | 2 | 656.34 | 122.58 | NR | U | N/A
|  | C14 Alkane | 18.337 | 277.14 | 1 | 2 | 554.28 | 120.82 | NR | U | N/A
|  | Alkene | 18.474 | 5.42 | 1 | 2 | 10.84 | 1.82 | NR | U | N/A
|  | Alkene | 18.554 | 5.47 | 1 | 2 | 10.94 | 1.18 | NR | U | N/A
|  | Branched C12 Alkene | 18.645 | 6.77 | 1 | 2 | 13.54 | 2.3 | NR | U | N/A
|  | Alkene | 18.748 | 28.91 | 1 | 2 | 57.82 | 8.48 | NR | U | N/A
|  | Alkene | 18.908 | 97.23 | 1 | 2 | 194.46 | 53.7 | NR | U | N/A
|  | octamethyl-Cyclotetrasiloxane | 20.474 | 10.33 | 1 | 2 | 20.66 | 20.66 | NR | TIC | 556-67-2
|  | 7-ethyl-2-methyl-4-Undecanol | 21.412 | 978.83 | 1 | 2 | 1957.66 | 339.56 | NR | MPC | 103-20-8
|  | C16-C20 Alkyl Benzene | 27.766 | 59.82 | 1 | 2 | 119.64 | 13.6 | NR | U | N/A
|  | Silane | 30.658 | 97.34 | 1 | 2 | 194.68 | 145.66 | NR | U | N/A
|  | C20-C25 Phthalate | 31.755 | 136.93 | 1 | 2 | 273.86 | 166.66 | NR | U | N/A
|  | Unknown | 33.07 | 163.09 | 1 | 2 | 326.18 | 105.64 | NR | U | N/A
|  | Unknown | 34.636 | 28.16 | 1 | 2 | 56.32 | 1.74 | NR | U | N/A
|  | Phthalate | 38.419 | 17.13 | 1 | 2 | 34.26 | 34.26 | NR | U | N/A
| Syringe S3 | Benzyl Alcohol | 13.388 | 2914.12 | 1 | 2 | 5828.24 | 767.9 | NR | IC | 100-51-6
|  | decamethyl-Cyclopentasiloxane | 14.999 | 9.19 | 1 | 2 | 18.38 | -28.32 | NR | MPC | 541-02-6
|  | C14 Alkene | 17.64 | 17.52 | 1 | 2 | 35.04 | 5.5 | NR | U | N/A
|  | C14 Alkene | 17.765 | 65.06 | 1 | 2 | 130.12 | 5.96 | NR | U | N/A
|  | C14 Alkene | 17.845 | 176.62 | 1 | 2 | 353.24 | 66.8 | NR | U | N/A
|  | C14 Alkene | 18.143 | 383.65 | 1 | 2 | 767.3 | 61.86 | NR | U | N/A
|  | C14 Alkane | 18.234 | 298.08 | 1 | 2 | 596.16 | 62.4 | NR | U | N/A
|  | C14 Alkane | 18.326 | 253.06 | 1 | 2 | 506.12 | 72.66 | NR | U | N/A
|  | Alkene | 18.474 | 5.11 | 1 | 2 | 10.22 | 1.2 | NR | U | N/A
|  | Alkene | 18.554 | 4.61 | 1 | 2 | 9.22 | -0.54 | NR | U | N/A
|  | Branched C12 Alkene | 18.645 | 4.91 | 1 | 2 | 9.82 | -1.42 | NR | U | N/A
|  | Alkene | 18.748 | 21.97 | 1 | 2 | 43.94 | -5.4 | NR | U | N/A
|  | Alkene | 18.908 | 72.88 | 1 | 2 | 145.76 | 5 | NR | U | N/A
|  | octamethyl-Cyclotetrasiloxane | 20.474 | 11.64 | 1 | 2 | 23.28 | 23.28 | NR | TIC | 556-67-2
|  | 7-ethyl-2-methyl-4-Undecanol | 21.412 | 1093.26 | 1 | 2 | 2186.52 | 568.42 | NR | MPC | 103-20-8
|  | C16-C20 Alkyl Benzene | 27.755 | 32.22 | 1 | 2 | 64.44 | -41.6 | NR | U | N/A
|  | Silane | 30.635 | 99.68 | 1 | 2 | 199.36 | 150.34 | NR | U | N/A
|  | C20-C25 Phthalate | 31.744 | 136.7 | 1 | 2 | 273.4 | 166.2 | NR | U | N/A
|  | Unknown | 33.07 | 150.28 | 1 | 2 | 300.56 | 80.02 | NR | U | N/A
|  | Unknown | 34.636 | 24.75 | 1 | 2 | 49.5 | -5.08 | NR | U | N/A
|  | Phthalate | 38.407 | 18.86 | 1 | 2 | 37.72 | 37.72 | NR | U | N/A
16-00866-N1 |  | SVOC
[8 empty] | Blank | Reportable
|  |  |  | Instrument | Concentration | Dilution | Corrected | Corrected | Not Reportable
|  | Compound | Ret Time | Result (ug/mL) | Factor | Factor | Result (ug/mL) | Result (ug/mL) | (R/NR) | ID | CAS #
3% S.A. | mb1603050OE | 3-Hexen-2-one | 7.547 | 1.3 | 1 | 2 | 2.6 | 2.6 | NR | MPC | 763-93-9
30 minute |  | hexamethyl-Cyclotrisiloxane | 7.901 | 10.38 | 1 | 2 | 20.76 | 20.76 | R | MPC | 541-05-9
|  | 2-(2-ethoxyethoxy)-Ethanol | 12.496 | 2.77 | 1 | 2 | 5.54 | 5.54 | NR | MPC | 111-90-0
|  | Benzyl Alcohol | 13.376 | 1728.02 | 1 | 2 | 3456.04 | 3456.04 | R | IC | 100-51-6
|  | Levoglucosenone | 14.908 | 18.51 | 1 | 2 | 37.02 | 37.02 | R | MPC | 37112-31-5
|  | decamethyl-Cyclopentasiloxane | 15.011 | 11.67 | 1 | 2 | 23.34 | 23.34 | R | MPC | 541-02-6
|  | C14 Alkene | 17.651 | 54.32 | 1 | 2 | 108.64 | 108.64 | R | U | N/A
|  | C14 Alkene | 17.777 | 195.92 | 1 | 2 | 391.84 | 391.84 | R | U | N/A
|  | C14 Alkene | 17.868 | 254.68 | 1 | 2 | 509.36 | 509.36 | R | U | N/A
|  | C14 Alkene | 18.017 | 5.82 | 1 | 2 | 11.64 | 11.64 | NR | U | N/A
|  | C14 Alkene | 18.165 | 557.86 | 1 | 2 | 1115.72 | 1115.72 | R | U | N/A
|  | C14 Alkane | 18.257 | 435.57 | 1 | 2 | 871.14 | 871.14 | R | U | N/A
|  | C14 Alkane | 18.348 | 368.07 | 1 | 2 | 736.14 | 736.14 | R | U | N/A
|  | Alkene | 18.485 | 21.73 | 1 | 2 | 43.46 | 43.46 | R | U | N/A
|  | Alkene | 18.565 | 22.05 | 1 | 2 | 44.1 | 44.1 | R | U | N/A
|  | Branched C12 Alkene | 18.645 | 28.04 | 1 | 2 | 56.08 | 56.08 | R | U | N/A
|  | Branched C12 Alkene | 18.703 | 20.56 | 1 | 2 | 41.12 | 41.12 | R | U | N/A
|  | Alkene | 18.76 | 74.35 | 1 | 2 | 148.7 | 148.7 | R | U | N/A
|  | Alkene | 18.828 | 8.01 | 1 | 2 | 16.02 | 16.02 | NR | U | N/A
|  | Alkene | 18.92 | 219.02 | 1 | 2 | 438.04 | 438.04 | R | U | N/A
|  | Alkene | 19.034 | 9.26 | 1 | 2 | 18.52 | 18.52 | NR | U | N/A
|  | octamethyl-Cyclotetrasiloxane | 20.417 | 1.47 | 1 | 2 | 2.94 | 2.94 | NR | TIC | 556-67-2
|  | Alkene | 21.24 | 12.59 | 1 | 2 | 25.18 | 25.18 | R | U | N/A
|  | 7-ethyl-2-methyl-4-Undecanol | 21.411 | 707.62 | 1 | 2 | 1415.24 | 1415.24 | R | MPC | 103-20-8
|  | 1,2-Dodecanediol | 21.777 | 6.91 | 1 | 2 | 13.82 | 13.82 | NR | TIC | 1119-87-5
|  | 9,10-dihydro-Anthracene | 24.817 | 16.08 | 1 | 2 | 32.16 | 32.16 | R | MPC | 613-31-0
|  | Alkyl Benzene | 27.561 | 13.1 | 1 | 2 | 26.2 | 26.2 | R | U | N/A
|  | C16-C20 Alkyl Benzene | 27.766 | 126.86 | 1 | 2 | 253.72 | 253.72 | R | U | N/A
|  | Unknown | 28.452 | 9.72 | 1 | 2 | 19.44 | 19.44 | NR | U | N/A
|  | Silane | 30.658 | 18.88 | 1 | 2 | 37.76 | 37.76 | R | U | N/A
|  | C20-C25 Phthalate | 31.755 | 26.97 | 1 | 2 | 53.94 | 53.94 | R | U | N/A
|  | Unknown | 33.081 | 38.27 | 1 | 2 | 76.54 | 76.54 | R | U | N/A
|  | Unknown | 34.533 | 8.53 | 1 | 2 | 17.06 | 17.06 | NR | U | N/A
| Control | 3-Hexen-2-one | 7.547 | 3.1 | 1 | 2 | 6.2 | 3.6 | NR | MPC | 763-93-9
|  | hexamethyl-Cyclotrisiloxane | 7.902 | 5.73 | 1 | 2 | 11.46 | -9.3 | NR | MPC | 541-05-9
|  | 2-(2-ethoxyethoxy)-Ethanol | 12.485 | 3.82 | 1 | 2 | 7.64 | -13.12 | NR | MPC | 111-90-0
|  | Benzyl Alcohol | 13.365 | 1842.73 | 1 | 2 | 3685.46 | 229.42 | NR | IC | 100-51-6
|  | Levoglucosenone | 14.908 | 10.61 | 1 | 2 | 21.22 | -15.8 | NR | MPC | 37112-31-5
|  | decamethyl-Cyclopentasiloxane | 14.999 | 16.35 | 1 | 2 | 32.7 | 9.36 | NR | MPC | 541-02-6
|  | C14 Alkene | 17.651 | 61.88 | 1 | 2 | 123.76 | 15.12 | NR | U | N/A
|  | C14 Alkene | 17.777 | 218.39 | 1 | 2 | 436.78 | 44.94 | NR | U | N/A
|  | C14 Alkene | 17.868 | 279.36 | 1 | 2 | 558.72 | 49.36 | NR | U | N/A
|  | C14 Alkene | 18.017 | 6.79 | 1 | 2 | 13.58 | 1.94 | NR | U | N/A
|  | C14 Alkene | 18.166 | 608.93 | 1 | 2 | 1217.86 | 102.14 | NR | U | N/A
|  | C14 Alkane | 18.257 | 474.2 | 1 | 2 | 948.4 | 77.26 | NR | U | N/A
|  | C14 Alkane | 18.348 | 404.97 | 1 | 2 | 809.94 | 73.8 | NR | U | N/A
|  | Alkene | 18.485 | 25.62 | 1 | 2 | 51.24 | 7.78 | NR | U | N/A
|  | Alkene | 18.566 | 26.37 | 1 | 2 | 52.74 | 8.64 | NR | U | N/A
|  | Branched C12 Alkene | 18.645 | 32.05 | 1 | 2 | 64.1 | 8.02 | NR | U | N/A
|  | Branched C12 Alkene | 18.703 | 22.76 | 1 | 2 | 45.52 | 4.4 | NR | U | N/A
|  | Alkene | 18.76 | 76.97 | 1 | 2 | 153.94 | 5.24 | NR | U | N/A
|  | Alkene | 18.828 | 9.06 | 1 | 2 | 18.12 | 2.1 | NR | U | N/A
|  | Alkene | 18.92 | 242.51 | 1 | 2 | 485.02 | 46.98 | NR | U | N/A
|  | Alkene | 19.034 | 12.6 | 1 | 2 | 25.2 | 6.68 | NR | U | N/A
|  | octamethyl-Cyclotetrasiloxane | 20.417 | 11.02 | 1 | 2 | 22.04 | 19.1 | NR | TIC | 556-67-2
|  | 7-ethyl-2-methyl-4-Undecanol | 21.389 | 668.97 | 1 | 2 | 1337.94 | -77.3 | NR | MPC | 103-20-8
|  | Unknown | 21.686 | 0.6 | 1 | 2 | 1.2 | 1.2 | NR | U | N/A
|  | 1,2-Dodecanediol | 21.777 | 4.48 | 1 | 2 | 8.96 | -4.86 | NR | TIC | 1119-87-5
|  | 9,10-dihydro-Anthracene | 24.818 | 9.79 | 1 | 2 | 19.58 | -12.58 | NR | MPC | 613-31-0
|  | Alkyl Benzene | 27.549 | 6.07 | 1 | 2 | 12.14 | -14.06 | NR | U | N/A
|  | C16-C20 Alkyl Benzene | 27.766 | 24.2 | 1 | 2 | 48.4 | -205.32 | NR | U | N/A
|  | Unknown | 27.835 | 8.16 | 1 | 2 | 16.32 | 16.32 | NR | U | N/A
|  | Unknown | 28.452 | 7.03 | 1 | 2 | 14.06 | -5.38 | NR | U | N/A
|  | Silane | 30.647 | 22.32 | 1 | 2 | 44.64 | 6.88 | NR | U | N/A
|  | C20-C25 Phthalate | 31.755 | 54.45 | 1 | 2 | 108.9 | 54.96 | NR | U | N/A
|  | Unknown | 33.081 | 131.97 | 1 | 2 | 263.94 | 187.4 | NR | U | N/A
|  | Unknown | 34.658 | 18.47 | 1 | 2 | 36.94 | 36.94 | R | U | N/A
|  | Unknown | 38.453 | 2.9 | 1 | 2 | 5.8 | 5.8 | NR | U | N/A
|  | Unknown | 40.728 | 2.44 | 1 | 2 | 4.88 | 4.88 | NR | U | N/A
| Catheter/Syringe S1 | Benzyl Alcohol | 13.376 | 1928.49 | 1 | 2 | 3856.98 | 400.94 | NR | IC | 100-51-6
|  | decamethyl-Cyclopentasiloxane | 15.011 | 11.09 | 1 | 2 | 22.18 | -1.16 | NR | MPC | 541-02-6
|  | C14 Alkene | 17.651 | 64.65 | 1 | 2 | 129.3 | 20.66 | NR | U | N/A
|  | C14 Alkene | 17.777 | 218.54 | 1 | 2 | 437.08 | 45.24 | NR | U | N/A
|  | C14 Alkene | 17.868 | 276.69 | 1 | 2 | 553.38 | 44.02 | NR | U | N/A
|  | C14 Alkene | 18.165 | 581.27 | 1 | 2 | 1162.54 | 46.82 | NR | U | N/A
|  | C14 Alkane | 18.257 | 454.9 | 1 | 2 | 909.8 | 38.66 | NR | U | N/A
|  | C14 Alkane | 18.348 | 396.44 | 1 | 2 | 792.88 | 56.74 | NR | U | N/A
|  | Alkene | 18.485 | 27.71 | 1 | 2 | 55.42 | 11.96 | NR | U | N/A
|  | Alkene | 18.566 | 29.98 | 1 | 2 | 59.96 | 15.86 | NR | U | N/A
|  | Branched C12 Alkene | 18.645 | 36.13 | 1 | 2 | 72.26 | 16.18 | NR | U | N/A
|  | Branched C12 Alkene | 18.703 | 25.79 | 1 | 2 | 51.58 | 10.46 | NR | U | N/A
|  | Alkene | 18.76 | 85.05 | 1 | 2 | 170.1 | 21.4 | NR | U | N/A
|  | Alkene | 18.828 | 10.61 | 1 | 2 | 21.22 | 5.2 | NR | U | N/A
|  | Alkene | 18.92 | 253.73 | 1 | 2 | 507.46 | 69.42 | NR | U | N/A
|  | Alkene | 19.034 | 14.71 | 1 | 2 | 29.42 | 10.9 | NR | U | N/A
|  | octamethyl-Cyclotetrasiloxane | 20.429 | 7.06 | 1 | 2 | 14.12 | 11.18 | NR | TIC | 556-67-2
|  | 7-ethyl-2-methyl-4-Undecanol | 21.4 | 672.05 | 1 | 2 | 1344.1 | -71.14 | NR | MPC | 103-20-8
|  | 9,10-dihydro-Anthracene | 24.817 | 19.02 | 1 | 2 | 38.04 | 5.88 | NR | MPC | 613-31-0
|  | Alkyl Benzene | 27.549 | 12.88 | 1 | 2 | 25.76 | -0.44 | NR | U | N/A
|  | C16-C20 Alkyl Benzene | 27.766 | 82.35 | 1 | 2 | 164.7 | -89.02 | NR | U | N/A
|  | Unknown | 27.835 | 10.14 | 1 | 2 | 20.28 | 20.28 | NR | U | N/A
|  | Unknown | 28.452 | 9.29 | 1 | 2 | 18.58 | -0.86 | NR | U | N/A
|  | Silane | 30.647 | 44.73 | 1 | 2 | 89.46 | 51.7 | NR | U | N/A
|  | C20-C25 Phthalate | 31.755 | 102.36 | 1 | 2 | 204.72 | 150.78 | NR | U | N/A
|  | Unknown | 33.081 | 200.18 | 1 | 2 | 400.36 | 323.82 | NR | U | N/A
|  | Unknown | 34.658 | 51.21 | 1 | 2 | 102.42 | 102.42 | NR | U | N/A
| Catheter/Syringe S2 | Benzyl Alcohol | 13.376 | 1882.8 | 1 | 2 | 3765.6 | 309.56 | NR | IC | 100-51-6
|  | decamethyl-Cyclopentasiloxane | 15.011 | 12.08 | 1 | 2 | 24.16 | 0.82 | NR | MPC | 541-02-6
|  | C14 Alkene | 17.651 | 70.97 | 1 | 2 | 141.94 | 33.3 | NR | U | N/A
|  | C14 Alkene | 17.777 | 233.56 | 1 | 2 | 467.12 | 75.28 | NR | U | N/A
|  | C14 Alkene | 17.868 | 285.82 | 1 | 2 | 571.64 | 62.28 | NR | U | N/A
|  | C14 Alkene | 18.177 | 598.79 | 1 | 2 | 1197.58 | 81.86 | NR | U | N/A
|  | C14 Alkane | 18.268 | 471.26 | 1 | 2 | 942.52 | 71.38 | NR | U | N/A
|  | C14 Alkane | 18.36 | 411.57 | 1 | 2 | 823.14 | 87 | NR | U | N/A
|  | Alkene | 18.485 | 38.4 | 1 | 2 | 76.8 | 33.34 | NR | U | N/A
|  | Alkene | 18.565 | 37.09 | 1 | 2 | 74.18 | 30.08 | NR | U | N/A
|  | Branched C12 Alkene | 18.657 | 41.09 | 1 | 2 | 82.18 | 26.1 | NR | U | N/A
|  | Branched C12 Alkene | 18.703 | 30.81 | 1 | 2 | 61.62 | 20.5 | NR | U | N/A
|  | Alkene | 18.76 | 94.41 | 1 | 2 | 188.82 | 40.12 | NR | U | N/A
|  | Alkene | 18.828 | 12 | 1 | 2 | 24 | 7.98 | NR | U | N/A
|  | Alkene | 18.92 | 281.48 | 1 | 2 | 562.96 | 124.92 | NR | U | N/A
|  | Alkene | 19.045 | 15.46 | 1 | 2 | 30.92 | 12.4 | NR | U | N/A
|  | octamethyl-Cyclotetrasiloxane | 20.428 | 6.39 | 1 | 2 | 12.78 | 9.84 | NR | TIC | 556-67-2
|  | 7-ethyl-2-methyl-4-Undecanol | 21.4 | 843.22 | 1 | 2 | 1686.44 | 271.2 | NR | MPC | 103-20-8
|  | 9,10-dihydro-Anthracene | 24.817 | 14.62 | 1 | 2 | 29.24 | -2.92 | NR | MPC | 613-31-0
|  | Alkyl Benzene | 27.549 | 10.8 | 1 | 2 | 21.6 | -4.6 | NR | U | N/A
|  | C16-C20 Alkyl Benzene | 27.766 | 64.76 | 1 | 2 | 129.52 | -124.2 | NR | U | N/A
|  | Unknown | 27.835 | 10.24 | 1 | 2 | 20.48 | 20.48 | NR | U | N/A
|  | Unknown | 28.452 | 8 | 1 | 2 | 16 | -3.44 | NR | U | N/A
|  | Silane | 30.647 | 30.67 | 1 | 2 | 61.34 | 23.58 | NR | U | N/A
|  | C20-C25 Phthalate | 31.755 | 66.07 | 1 | 2 | 132.14 | 78.2 | NR | U | N/A
|  | Unknown | 33.081 | 104.02 | 1 | 2 | 208.04 | 131.5 | NR | U | N/A
|  | Unknown | 34.658 | 34.81 | 1 | 2 | 69.62 | 69.62 | NR | U | N/A
| Catheter/Syringe S3 | Benzyl Alcohol | 13.399 | 2333.09 | 1 | 2 | 4666.18 | 1210.14 | NR | IC | 100-51-6
|  | decamethyl-Cyclopentasiloxane | 15.011 | 17.15 | 1 | 2 | 34.3 | 10.96 | NR | MPC | 541-02-6
|  | C14 Alkene | 17.651 | 28.07 | 1 | 2 | 56.14 | -52.5 | NR | U | N/A
|  | C14 Alkene | 17.777 | 111.62 | 1 | 2 | 223.24 | -168.6 | NR | U | N/A
|  | C14 Alkene | 17.857 | 186.17 | 1 | 2 | 372.34 | -137.02 | NR | U | N/A
|  | C14 Alkene | 18.154 | 434.92 | 1 | 2 | 869.84 | -245.88 | NR | U | N/A
|  | C14 Alkane | 18.246 | 335.31 | 1 | 2 | 670.62 | -200.52 | NR | U | N/A
|  | C14 Alkane | 18.337 | 281.94 | 1 | 2 | 563.88 | -172.26 | NR | U | N/A
|  | Alkene | 18.485 | 11.62 | 1 | 2 | 23.24 | -20.22 | NR | U | N/A
|  | Alkene | 18.566 | 8.59 | 1 | 2 | 17.18 | -26.92 | NR | U | N/A
|  | Branched C12 Alkene | 18.645 | 8.12 | 1 | 2 | 16.24 | -39.84 | NR | U | N/A
|  | Branched C12 Alkene | 18.703 | 9.43 | 1 | 2 | 18.86 | -22.26 | NR | U | N/A
|  | Alkene | 18.76 | 33.97 | 1 | 2 | 67.94 | -80.76 | NR | U | N/A
|  | Alkene | 18.828 | 0.53 | 1 | 2 | 1.06 | -14.96 | NR | U | N/A
|  | Alkene | 18.92 | 139.28 | 1 | 2 | 278.56 | -159.48 | NR | U | N/A
|  | Alkene | 19.034 | 4.61 | 1 | 2 | 9.22 | -9.3 | NR | U | N/A
|  | octamethyl-Cyclotetrasiloxane | 20.429 | 13.53 | 1 | 2 | 27.06 | 24.12 | NR | TIC | 556-67-2
|  | 7-ethyl-2-methyl-4-Undecanol | 21.412 | 659.63 | 1 | 2 | 1319.26 | -95.98 | NR | MPC | 103-20-8
|  | 9,10-dihydro-Anthracene | 24.818 | 7.56 | 1 | 2 | 15.12 | -17.04 | NR | MPC | 613-31-0
|  | Alkyl Benzene | 27.561 | 3.27 | 1 | 2 | 6.54 | -19.66 | NR | U | N/A
|  | C16-C20 Alkyl Benzene | 27.766 | 13.4 | 1 | 2 | 26.8 | -226.92 | NR | U | N/A
|  | Unknown | 27.835 | 1.29 | 1 | 2 | 2.58 | 2.58 | NR | U | N/A
|  | Unknown | 28.452 | 0.98 | 1 | 2 | 1.96 | -17.48 | NR | U | N/A
|  | Silane | 30.647 | 45.24 | 1 | 2 | 90.48 | 52.72 | NR | U | N/A
|  | C20-C25 Phthalate | 31.755 | 125.15 | 1 | 2 | 250.3 | 196.36 | NR | U | N/A
|  | Unknown | 33.081 | 288.27 | 1 | 2 | 576.54 | 500 | NR | U | N/A
|  | Unknown | 34.647 | 72.57 | 1 | 2 | 145.14 | 145.14 | NR | U | N/A
| Syringe S1 | 3-Hexen-2-one | 7.547 | 9.5 | 1 | 2 | 19 | 16.4 | NR | MPC | 763-93-9
|  | hexamethyl-Cyclotrisiloxane | 7.901 | 13.61 | 1 | 2 | 27.22 | 6.46 | NR | MPC | 541-05-9
|  | 2-(2-ethoxyethoxy)-Ethanol | 12.508 | 9.68 | 1 | 2 | 19.36 | -1.4 | NR | MPC | 111-90-0
|  | Benzyl Alcohol | 13.388 | 2109.41 | 1 | 2 | 4218.82 | 762.78 | NR | IC | 100-51-6
|  | decamethyl-Cyclopentasiloxane | 15.011 | 17.29 | 1 | 2 | 34.58 | 11.24 | NR | MPC | 541-02-6
|  | C14 Alkene | 17.651 | 64.35 | 1 | 2 | 128.7 | 20.06 | NR | U | N/A
|  | C14 Alkene | 17.777 | 221.29 | 1 | 2 | 442.58 | 50.74 | NR | U | N/A
|  | C14 Alkene | 17.868 | 288.88 | 1 | 2 | 577.76 | 577.76 | NR | U | N/A
|  | C14 Alkene | 18.017 | 7.52 | 1 | 2 | 15.04 | 15.04 | NR | U | N/A
|  | C14 Alkene | 18.165 | 611 | 1 | 2 | 1222 | 106.28 | NR | U | N/A
|  | C14 Alkane | 18.257 | 474.68 | 1 | 2 | 949.36 | 78.22 | NR | U | N/A
|  | C14 Alkane | 18.348 | 410.95 | 1 | 2 | 821.9 | 85.76 | NR | U | N/A
|  | Alkene | 18.485 | 29.02 | 1 | 2 | 58.04 | 14.58 | NR | U | N/A
|  | Alkene | 18.565 | 31.8 | 1 | 2 | 63.6 | 19.5 | NR | U | N/A
|  | Branched C12 Alkene | 18.645 | 34.86 | 1 | 2 | 69.72 | 13.64 | NR | U | N/A
|  | Branched C12 Alkene | 18.703 | 25.81 | 1 | 2 | 51.62 | 10.5 | NR | U | N/A
|  | Alkene | 18.76 | 86.56 | 1 | 2 | 173.12 | 24.42 | NR | U | N/A
|  | Alkene | 18.92 | 257.51 | 1 | 2 | 515.02 | 76.98 | NR | U | N/A
|  | Alkene | 19.034 | 12.07 | 1 | 2 | 24.14 | 5.62 | NR | U | N/A
|  | octamethyl-Cyclotetrasiloxane | 20.428 | 13.23 | 1 | 2 | 26.46 | 23.52 | NR | TIC | 556-67-2
|  | 7-ethyl-2-methyl-4-Undecanol | 21.4 | 882.4 | 1 | 2 | 1764.8 | 349.56 | NR | MPC | 103-20-8
|  | Unknown | 21.686 | 2.33 | 1 | 2 | 4.66 | 4.66 | NR | U | N/A
|  | 1,2-Dodecanediol | 21.777 | 3.39 | 1 | 2 | 6.78 | -7.04 | NR | TIC | 1119-87-5
|  | 9,10-dihydro-Anthracene | 24.817 | 18.04 | 1 | 2 | 36.08 | 3.92 | NR | MPC | 613-31-0
|  | Alkyl Benzene | 27.549 | 9.6 | 1 | 2 | 19.2 | -7 | NR | U | N/A
|  | C16-C20 Alkyl Benzene | 27.766 | 38.85 | 1 | 2 | 77.7 | -176.02 | NR | U | N/A
|  | Silane | 30.647 | 52.08 | 1 | 2 | 104.16 | 66.4 | NR | U | N/A
|  | C20-C25 Phthalate | 31.755 | 108.76 | 1 | 2 | 217.52 | 163.58 | NR | U | N/A
|  | Unknown | 33.081 | 249.94 | 1 | 2 | 499.88 | 423.34 | NR | U | N/A
|  | Unknown | 34.658 | 71.52 | 1 | 2 | 143.04 | 143.04 | NR | U | N/A
|  | Unknown | 38.442 | 11.08 | 1 | 2 | 22.16 | 22.16 | NR | U | N/A
|  | Unknown | 40.728 | 6.44 | 1 | 2 | 12.88 | 12.88 | NR | U | N/A
| Syringe S2 | 3-Hexen-2-one | 7.547 | 7.12 | 1 | 2 | 14.24 | 11.64 | NR | MPC | 763-93-9
|  | hexamethyl-Cyclotrisiloxane | 7.902 | 8.94 | 1 | 2 | 17.88 | -2.88 | NR | MPC | 541-05-9
|  | 2-(2-ethoxyethoxy)-Ethanol | 12.531 | 10.93 | 1 | 2 | 21.86 | 1.1 | NR | MPC | 111-90-0
|  | Benzyl Alcohol | 13.411 | 2540.41 | 1 | 2 | 5080.82 | 1624.78 | NR | IC | 100-51-6
|  | decamethyl-Cyclopentasiloxane | 15.011 | 9.48 | 1 | 2 | 18.96 | -4.38 | NR | MPC | 541-02-6
|  | C14 Alkene | 17.651 | 33.6 | 1 | 2 | 67.2 | -41.44 | NR | U | N/A
|  | C14 Alkene | 17.777 | 120.38 | 1 | 2 | 240.76 | -151.08 | NR | U | N/A
|  | C14 Alkene | 17.857 | 223.35 | 1 | 2 | 446.7 | 446.7 | NR | U | N/A
|  | C14 Alkene | 18.166 | 479.91 | 1 | 2 | 959.82 | -155.9 | NR | U | N/A
|  | C14 Alkane | 18.257 | 363.83 | 1 | 2 | 727.66 | -143.48 | NR | U | N/A
|  | C14 Alkane | 18.348 | 303.17 | 1 | 2 | 606.34 | -129.8 | NR | U | N/A
|  | Alkene | 18.485 | 9.95 | 1 | 2 | 19.9 | -23.56 | NR | U | N/A
|  | Alkene | 18.566 | 8.15 | 1 | 2 | 16.3 | -27.8 | NR | U | N/A
|  | Branched C12 Alkene | 18.657 | 8.7 | 1 | 2 | 17.4 | -38.68 | NR | U | N/A
|  | Branched C12 Alkene | 18.703 | 9.45 | 1 | 2 | 18.9 | -22.22 | NR | U | N/A
|  | Alkene | 18.76 | 41.03 | 1 | 2 | 82.06 | -66.64 | NR | U | N/A
|  | Alkene | 18.92 | 152.29 | 1 | 2 | 304.58 | -133.46 | NR | U | N/A
|  | Alkene | 19.034 | 5.61 | 1 | 2 | 11.22 | -7.3 | NR | U | N/A
|  | octamethyl-Cyclotetrasiloxane | 20.463 | 16.21 | 1 | 2 | 32.42 | 29.48 | NR | TIC | 556-67-2
|  | 7-ethyl-2-methyl-4-Undecanol | 21.423 | 953.74 | 1 | 2 | 1907.48 | 492.24 | NR | MPC | 103-20-8
|  | Unknown | 21.686 | 2.67 | 1 | 2 | 5.34 | 5.34 | NR | U | N/A
|  | 1,2-Dodecanediol | 21.789 | 6.28 | 1 | 2 | 12.56 | -1.26 | NR | TIC | 1119-87-5
|  | 9,10-dihydro-Anthracene | 24.818 | 7.44 | 1 | 2 | 14.88 | -17.28 | NR | MPC | 613-31-0
|  | Alkyl Benzene | 27.561 | 4.75 | 1 | 2 | 9.5 | -16.7 | NR | U | N/A
|  | C16-C20 Alkyl Benzene | 27.766 | 43.56 | 1 | 2 | 87.12 | -166.6 | NR | U | N/A
|  | Silane | 30.647 | 53.41 | 1 | 2 | 106.82 | 69.06 | NR | U | N/A
|  | Phthalate | 31.641 | 2.85 | 1 | 2 | 5.7 | 5.7 | NR | U | N/A
|  | C20-C25 Phthalate | 31.755 | 106.17 | 1 | 2 | 212.34 | 158.4 | NR | U | N/A
|  | Unknown | 33.081 | 283.33 | 1 | 2 | 566.66 | 490.12 | NR | U | N/A
|  | Unknown | 34.658 | 89.52 | 1 | 2 | 179.04 | 179.04 | NR | U | N/A
|  | Unknown | 38.442 | 15.58 | 1 | 2 | 31.16 | 31.16 | NR | U | N/A
|  | Unknown | 40.728 | 8.82 | 1 | 2 | 17.64 | 17.64 | NR | U | N/A
| Syringe S3 | 3-Hexen-2-one | 7.547 | 11.24 | 1 | 2 | 22.48 | 19.88 | NR | MPC | 763-93-9
|  | hexamethyl-Cyclotrisiloxane | 7.902 | 10.33 | 1 | 2 | 20.66 | -0.1 | NR | MPC | 541-05-9
|  | 2-(2-ethoxyethoxy)-Ethanol | 12.508 | 12.83 | 1 | 2 | 25.66 | 4.9 | NR | MPC | 111-90-0
|  | Benzyl Alcohol | 13.388 | 2319.31 | 1 | 2 | 4638.62 | 1182.58 | NR | IC | 100-51-6
|  | decamethyl-Cyclopentasiloxane | 15.011 | 11.12 | 1 | 2 | 22.24 | -1.1 | NR | MPC | 541-02-6
|  | C14 Alkene | 17.651 | 54.78 | 1 | 2 | 109.56 | 0.92 | NR | U | N/A
|  | C14 Alkene | 17.777 | 196.34 | 1 | 2 | 392.68 | 0.84 | NR | U | N/A
|  | C14 Alkene | 17.868 | 263.13 | 1 | 2 | 526.26 | 526.26 | NR | U | N/A
|  | C14 Alkene | 18.017 | 8.27 | 1 | 2 | 16.54 | 16.54 | NR | U | N/A
|  | C14 Alkene | 18.165 | 589.81 | 1 | 2 | 1179.62 | 63.9 | NR | U | N/A
|  | C14 Alkane | 18.257 | 462.73 | 1 | 2 | 925.46 | 54.32 | NR | U | N/A
|  | C14 Alkane | 18.348 | 396.33 | 1 | 2 | 792.66 | 56.52 | NR | U | N/A
|  | Alkene | 18.485 | 29.58 | 1 | 2 | 59.16 | 15.7 | NR | U | N/A
|  | Alkene | 18.565 | 29.54 | 1 | 2 | 59.08 | 14.98 | NR | U | N/A
|  | Branched C12 Alkene | 18.645 | 34.18 | 1 | 2 | 68.36 | 12.28 | NR | U | N/A
|  | Branched C12 Alkene | 18.703 | 23.41 | 1 | 2 | 46.82 | 5.7 | NR | U | N/A
|  | Alkene | 18.76 | 79.35 | 1 | 2 | 158.7 | 10 | NR | U | N/A
|  | Alkene | 18.92 | 236.04 | 1 | 2 | 472.08 | 34.04 | NR | U | N/A
|  | Alkene | 19.034 | 11.77 | 1 | 2 | 23.54 | 5.02 | NR | U | N/A
|  | octamethyl-Cyclotetrasiloxane | 20.451 | 10.52 | 1 | 2 | 21.04 | 18.1 | NR | TIC | 556-67-2
|  | 7-ethyl-2-methyl-4-Undecanol | 21.4 | 906.86 | 1 | 2 | 1813.72 | 398.48 | NR | MPC | 103-20-8
|  | Unknown | 21.686 | 10.75 | 1 | 2 | 21.5 | 21.5 | NR | U | N/A
|  | 1,2-Dodecanediol | 21.777 | 12.1 | 1 | 2 | 24.2 | 10.38 | NR | TIC | 1119-87-5
|  | 9,10-dihydro-Anthracene | 24.817 | 13.11 | 1 | 2 | 26.22 | -5.94 | NR | MPC | 613-31-0
|  | Alkyl Benzene | 27.561 | 10.1 | 1 | 2 | 20.2 | -6 | NR | U | N/A
|  | C16-C20 Alkyl Benzene | 27.766 | 83.88 | 1 | 2 | 167.76 | -85.96 | NR | U | N/A
|  | Silane | 30.647 | 61.74 | 1 | 2 | 123.48 | 85.72 | NR | U | N/A
|  | C20-C25 Phthalate | 31.755 | 115.64 | 1 | 2 | 231.28 | 177.34 | NR | U | N/A
|  | Unknown | 33.081 | 270.29 | 1 | 2 | 540.58 | 464.04 | NR | U | N/A
|  | Unknown | 34.658 | 93.11 | 1 | 2 | 186.22 | 186.22 | NR | U | N/A
|  | Unknown | 38.453 | 21.31 | 1 | 2 | 42.62 | 42.62 | NR | U | N/A
|  | Unknown | 40.728 | 11.89 | 1 | 2 | 23.78 | 23.78 | NR | U | N/A

## NVOC
16-00866-N1 |  | NVOC
Blank
|  |  |  | Instrument | Concentration | Dilution | Corrected | Corrected
|  | Compound | Ret Time | Result (ug/mL) | Factor | Factor | Result (ug/mL) | Result (ug/mL) | ID | CAS #
Water | Control | Compound | 6.12 | 10 | 5 | 1 | 2 | 2 | IC | 57-10-3
| Sample 1 | None | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A
| Sample 2 | None | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A
| Sample 3 | None | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A
BRL - Below reporting Limit (<1 ug/mL)
Compounds observed in blank not reported, as they are not test article/control related.
Corrected Resulted (ug/mL) = (Instrument Result (ug/mL) * Dilution Factor)/Concentration Factor
Blank/Control Corrected Result (ug/mL) = Correct Result (ug/mL) - Corrected Result Corrected for Blank/Control (ug/mL)
Concentration Factor = Initial Extraction Volume/Concentrated Volume
Highlighted results reported

## LCUV
16-00866-N1 |  | LC/UV
|  | Additional Peaks | Ret Time | Wavelenth(s) (nm)
Water | Control | N/A | N/A | N/A
| Sample 1 | None | N/A | N/A
| Sample 2 | None | N/A | N/A
| Sample 3 | None | N/A | N/A

## pH - Conductivity (hidden)
16-00866-N1 |  | pH
15 minute
|  | Result | Average | St dev
WFI | Control | 6.98 | 6.98 | N/A
| Sample 1 | 6.54 | 6.5833 | 0.29738
| Sample 2 | 6.31
| Sample 3 | 6.9
16-00866-N1 |  | Conductivity
15 minute |  |  | Average
|  | Result (uS/cm) | Result (uS/cm) | St dev
WFI | Control | 1.93 | 1.93 | N/A
| Sample 1 | 2.03 | 2.44 | 0.43715
| Sample 2 | 2.9
| Sample 3 | 2.39

## TOC (hidden)
16-00866-N1 |  | TOC
15 Minute |  | Instrument | Blank Correct
|  | Result (mg/L) | Result (mg/L) | Average (mg/L) | Std Dev
| Control | 0.2077 | 0.2077 | 0.2077 | N/A
| Sample 1 | 0.3719 | 0.1642 | 0.2602 | 0.083173
| Sample 2 | 0.5135 | 0.3058
| Sample 3 | 0.5183 | 0.3106
Correct Result (ug/mL) = Instrument Result (mg/L) - Control Result (mg/L)

## CAS #
VOC |  | SVOC |  | NVOC
Compound | CAS #
Acetone | 67-64-1 |  |  | Compound | CAS #
Caprolactam | 105-60-2 | Compound | CAS # | Irganox 245 | 36443-68-2
Tert butyl alcohol | 75-65-0 | Diethylphthalate | 84-66-2 | BHT | 128-37-0
|  | Dimethylphthalate | 131-11-3 | Myristic acid | 544-63-8
2 methyl 1 propene | 115-11-7 | bis (2-ethylhexyl) phthalate | 117-81-7 | Oleamide | 301-02-0
ethyl alcohol | 64-17-5 | Bisphenol A | 80-05-7 | A.O. 2246 | 119-47-1
2-ethyl 1 hexanol | 104-76-7 |  |  | Palmitic acid | 57-10-3
Butanal | 123-72-8 | Trioctyl Trimellitate (TOTM) | 3319-31-1 | Oleic acid | 112-80-1
Pentanal | 110-62-3 |  |  | Stearic acid | 57-11-4
Hexanal | 66-25-1 |  |  | Erucamide | 112-84-5
|  |  |  | Irganox 1010 | 98584-37-3
Benzyl alcohol | 100-51-6 |  |  | Irganox 1330 | 1709-70-5
2-butanone | 78-93-3 |  |  | Irganox 1076 | 2082-79-3
2-hexanone | 591-78-6 |  |  | Irgafos 168 | 31570-04-4
|  |  |  | Irgafos 126 | 26741-53-7
|  |  |  | Irganox 3114 | 27676-62-6
1-propene, 2methyl- | 115-11-7
methylene chloride | 75-09-2
acetonitrile | 75-05-8
chloromethane | 74-87-3
bromomethane | 74-83-9
IPA | 67-63-0
tetrahydrofuran | 109-99-9
acrylonitrile | 1424-48-2
styrene | 100-42-5
chloro benzene | 108-90-7
naphthalene | 91-20-3
4-methyl-2-pentanone | 108-10-1
ethylbenzene | 100-41-4
m,p-xylene | 1330-20-7
n-hexane | 110-54-3
toluene | 108-88-3
Benzene | 71-43-2
1,1,1-trichloroethane | 71-55-6
Carbon tetrachloride | 56-23-5
ethyl tert-butyl ether | 637-92-3
1,3,5-trimethylbenzene | 108-67-8
1,2,4-trimethylbenzene | 95-63-6
Chloroform | 67-66-3
2,2,4,6,6-pentamethylheptane | 13475-82-6